## Intro
Supplementary Material
Establishment and Thorough External Validation of a FTIR Spectroscopy Classifier for Salmonella Serogroup Differentiation
Authors: Helene Oberreuter1*, Miriam Cordovana2, Martin Dyk1, Jörg Rau1
1Chemical and Veterinary Analysis Agency Stuttgart (CVUAS), Fellbach, Germany
2Bruker Daltonics GmbH & Co. KG, Bremen, Germany
*Corresponding author: Helene Oberreuter, Chemical and Veterinary Analysis Agency Stuttgart (CVUAS), Schaflandstr. 3/2, 70736 Fellbach, Germany, helene.oberreuter@cvuas.bwl.de
Supplement 2
Excel file including three worksheets
- Worksheet IR BT target parm: Classification results of the targeted identification of Salmonella serogroup C1 (O:7) following the Guidelines for Validating Species Identifications using MALDI-ToF-MS pursuant to §64 of the German Food and Feed code (LFGB), coordinated by the Federal Office of Consumer Protection and Food Safety, Berlin, Germany (BVL) 28 Oct 2022: target parameter.
- Worksheet IR BT non-target parm: Classification results of the targeted identification of Salmonella serogroup C1 (O:7) following the Guidelines for Validating Species Identifications using MALDI-ToF-MS pursuant to §64 of the German Food and Feed code (LFGB), coordinated by the Federal Office of Consumer Protection and Food Safety, Berlin, Germany (BVL) 28 Oct 2022: non-target parameter.
- Report (cf. Fig 2): Summary report of the targeted identification of Salmonella serogroup C1 (O:7) following the Guidelines for Validating Species Identifications using MALDI-ToF-MSpursuant to §64 of the German Food and Feed code (LFGB), coordinated by the Federal Office of Consumer Protection and Food Safety, Berlin, Germany (BVL) 28 Oct 2022.

## IR BT target parm
Method |  | Salmonella O-group differentiation by Bruker salmonella o-groups classifier v3 | [7 empty] | target parameter | O:7 (C1-C4)
Differentation |  | External validation of the a.m. classifier; incubation on COL-SB 37°C 24h | isolates | 25 | [5 empty] | non-target parameter | non - O:7 (C1-C4)
Date of validation |  | see report
|  |  | spectra | 150
[5 empty] | classified | number | percentage |  |  | outlier score | 150
[5 empty] | correct | 147 | 98% |  |  | Number in green | 76
[5 empty] | questionable | 2 | 1.333% |  |  | Number in yellow | 72
[5 empty] | wrong | 1 | 0.6667% |  |  | Number in red | 2
[11 empty] | 147 | 2 | 1
Isolate |  | Name (incl. spot no.) | Isolate description | Isolate ID | QT result: wahr=true=pass; falsch=false=fail | Factual classification | IR BT classification | Classifi-cation correkt? (i.e. <> target parm) => 1 | irrelevant! classification score: if <0.6 => invalid | outlier score: <2,5: green, 2,5 - 3,5: yellow, >3,5: red | if outlier score </=3,5 AND classification = factual classification => 1 (correct) | if outlier score >3,5) [classification irrelevant] => 1 (questionable) | if outlier score </=3,5 AND classification <> factual classification => 1 (wrong)
# | 1 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 2323_2_B10.0 |  | CVUAS 2323.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.831 | 2.474 | 1 | 0 | 0
| 1 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 2323_2_B11.0 |  | CVUAS 2323.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.865 | 2.783 | 1 | 0 | 0
| 1 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 2323_2_B12.0 |  | CVUAS 2323.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.735 | 2.772 | 1 | 0 | 0
| 1 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 2323_2_B10.0 |  | CVUAS 2323.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.901 | 2.277 | 1 | 0 | 0
| 1 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 2323_2_B11.0 |  | CVUAS 2323.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.87 | 2.543 | 1 | 0 | 0
| 1 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 2323_2_B12.0 |  | CVUAS 2323.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.876 | 2.427 | 1 | 0 | 0
# | 2 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11506_3_H7.0 |  | CVUAS 11506.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.924 | 2.465 | 1 | 0 | 0
| 2 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11506_3_H8.0 |  | CVUAS 11506.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.943 | 2.388 | 1 | 0 | 0
| 2 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11506_3_H9.0 |  | CVUAS 11506.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.955 | 2.438 | 1 | 0 | 0
| 2 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11506_3_H7.0 |  | CVUAS 11506.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.8 | 2.818 | 1 | 0 | 0
| 2 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11506_3_H8.0 |  | CVUAS 11506.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.658 | 2.457 | 1 | 0 | 0
| 2 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11506_3_H9.0 |  | CVUAS 11506.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.711 | 2.465 | 1 | 0 | 0
# | 3 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31253_E4.0 |  | CVUAS 31253 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.98 | 2.477 | 1 | 0 | 0
| 3 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31253_E5.0 |  | CVUAS 31253 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.978 | 2.594 | 1 | 0 | 0
| 3 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31253_E6.0 |  | CVUAS 31253 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.979 | 2.609 | 1 | 0 | 0
| 3 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31253_E4.0 |  | CVUAS 31253 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.984 | 2.142 | 1 | 0 | 0
| 3 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31253_E5.0 |  | CVUAS 31253 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.983 | 2.121 | 1 | 0 | 0
| 3 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31253_E6.0 |  | CVUAS 31253 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.987 | 2.299 | 1 | 0 | 0
# | 4 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31297_H4.0 |  | CVUAS 31297 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.149 | 3.024 | 1 | 0 | 0
| 4 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31297_H5.0 |  | CVUAS 31297 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.127 | 3.233 | 1 | 0 | 0
| 4 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31297_H6.0 |  | CVUAS 31297 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.149 | 3.315 | 1 | 0 | 0
| 4 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31297_H4.0 |  | CVUAS 31297 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.513 | 2.42 | 1 | 0 | 0
| 4 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31297_H5.0 |  | CVUAS 31297 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.478 | 2.523 | 1 | 0 | 0
| 4 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31297_H6.0 |  | CVUAS 31297 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.486 | 2.612 | 1 | 0 | 0
# | 5 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32492_A10.0 |  | CVUAS 32492 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.926 | 2.348 | 1 | 0 | 0
| 5 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32492_A11.0 |  | CVUAS 32492 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.93 | 2.637 | 1 | 0 | 0
| 5 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32492_A12.0 |  | CVUAS 32492 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.844 | 3.063 | 1 | 0 | 0
| 5 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32492_A10.0 |  | CVUAS 32492 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.825 | 3.437 | 1 | 0 | 0
| 5 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32492_A11.0 |  | CVUAS 32492 | FALSE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.793 | 3.459 | 1 | 0 | 0
| 5 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32492_A12.0 |  | CVUAS 32492 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.904 | 3.108 | 1 | 0 | 0
# | 6 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32981_H10.0 |  | CVUAS 32981 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.936 | 2.515 | 1 | 0 | 0
| 6 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32981_H11.0 |  | CVUAS 32981 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.947 | 2.375 | 1 | 0 | 0
| 6 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32981_H12.0 |  | CVUAS 32981 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.872 | 2.831 | 1 | 0 | 0
| 6 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32981_H10.0 |  | CVUAS 32981 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.918 | 2.558 | 1 | 0 | 0
| 6 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32981_H11.0 |  | CVUAS 32981 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.898 | 2.466 | 1 | 0 | 0
| 6 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32981_H12.0 |  | CVUAS 32981 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.897 | 2.67 | 1 | 0 | 0
# | 7 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33075_A4.0 |  | CVUAS 33075 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.981 | 2.23 | 1 | 0 | 0
| 7 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33075_A5.0 |  | CVUAS 33075 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.978 | 2.432 | 1 | 0 | 0
| 7 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33075_A6.0 |  | CVUAS 33075 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.984 | 2.503 | 1 | 0 | 0
| 7 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33075_A4.0 |  | CVUAS 33075 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.913 | 2.091 | 1 | 0 | 0
| 7 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33075_A5.0 |  | CVUAS 33075 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.855 | 2.082 | 1 | 0 | 0
| 7 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33075_A6.0 |  | CVUAS 33075 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.786 | 2.242 | 1 | 0 | 0
# | 8 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 35791_H4.0 |  | CVUAS 35791 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.939 | 2.075 | 1 | 0 | 0
| 8 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 35791_H5.0 |  | CVUAS 35791 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.956 | 2.23 | 1 | 0 | 0
| 8 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 35791_H6.0 |  | CVUAS 35791 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.93 | 2.182 | 1 | 0 | 0
| 8 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 35791_H4.0 |  | CVUAS 35791 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.917 | 2.175 | 1 | 0 | 0
| 8 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 35791_H5.0 |  | CVUAS 35791 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.908 | 2.052 | 1 | 0 | 0
| 8 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 35791_H6.0 |  | CVUAS 35791 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.92 | 2.335 | 1 | 0 | 0
# | 9 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4808_3_F4.0 |  | CVUAS 4808.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.987 | 2.311 | 1 | 0 | 0
| 9 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4808_3_F5.0 |  | CVUAS 4808.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.991 | 2.498 | 1 | 0 | 0
| 9 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4808_3_F6.0 |  | CVUAS 4808.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.992 | 2.452 | 1 | 0 | 0
| 9 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4808_3_F4.0 |  | CVUAS 4808.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.918 | 2.029 | 1 | 0 | 0
| 9 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4808_3_F5.0 |  | CVUAS 4808.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.915 | 2.09 | 1 | 0 | 0
| 9 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4808_3_F6.0 |  | CVUAS 4808.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.919 | 2.147 | 1 | 0 | 0
# | 10 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5138_3_E7.0 |  | CVUAS 5138.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.983 | 2.684 | 1 | 0 | 0
| 10 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5138_3_E8.0 |  | CVUAS 5138.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.978 | 2.591 | 1 | 0 | 0
| 10 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5138_3_E9.0 |  | CVUAS 5138.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.973 | 2.679 | 1 | 0 | 0
| 10 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5138_3_E7.0 |  | CVUAS 5138.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.955 | 2.296 | 1 | 0 | 0
| 10 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5138_3_E8.0 |  | CVUAS 5138.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.967 | 2.642 | 1 | 0 | 0
| 10 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5138_3_E9.0 |  | CVUAS 5138.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.984 | 2.79 | 1 | 0 | 0
# | 11 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 85_2_A4.0 |  | CVUAS 85.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.797 | 2.04 | 1 | 0 | 0
| 11 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 85_2_A5.0 |  | CVUAS 85.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.861 | 2.042 | 1 | 0 | 0
| 11 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 85_2_A6.0 |  | CVUAS 85.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.733 | 2.238 | 1 | 0 | 0
| 11 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 85_2_A4.0 |  | CVUAS 85.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.875 | 1.873 | 1 | 0 | 0
| 11 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 85_2_A5.0 |  | CVUAS 85.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.884 | 1.944 | 1 | 0 | 0
| 11 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 85_2_A6.0 |  | CVUAS 85.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.897 | 2.175 | 1 | 0 | 0
# | 12 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7091_2_F4.0 |  | CVUAS 7091.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.959 | 2.509 | 1 | 0 | 0
| 12 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7091_2_F5.0 |  | CVUAS 7091.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.953 | 2.298 | 1 | 0 | 0
| 12 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7091_2_F6.0 |  | CVUAS 7091.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.956 | 2.529 | 1 | 0 | 0
| 12 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7091_2_F4.0 |  | CVUAS 7091.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.667 | 2.153 | 1 | 0 | 0
| 12 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7091_2_F5.0 |  | CVUAS 7091.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.699 | 2.281 | 1 | 0 | 0
| 12 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7091_2_F6.0 |  | CVUAS 7091.2 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.68 | 2.187 | 1 | 0 | 0
# | 13 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 326_4_D4.0 |  | CVUAS 326.4 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.19 | 2.394 | 1 | 0 | 0
| 13 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 326_4_D5.0 |  | CVUAS 326.4 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.181 | 2.246 | 1 | 0 | 0
| 13 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 326_4_D6.0 |  | CVUAS 326.4 | TRUE | O:7 (C1-C4) | O:6,14 (H) | 0 | 0.152 | 2.315 | 0 | 0 | 1
| 13 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 326_4_D4.0 |  | CVUAS 326.4 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.854 | 2.519 | 1 | 0 | 0
| 13 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 326_4_D5.0 |  | CVUAS 326.4 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.392 | 2.352 | 1 | 0 | 0
| 13 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 326_4_D6.0 |  | CVUAS 326.4 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.361 | 2.277 | 1 | 0 | 0
# | 14 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35069_A7.0 |  | CVUAS 35069 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.95 | 2.911 | 1 | 0 | 0
| 14 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35069_A8.0 |  | CVUAS 35069 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.965 | 2.942 | 1 | 0 | 0
| 14 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35069_A9.0 |  | CVUAS 35069 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.942 | 2.871 | 1 | 0 | 0
| 14 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35069_A7.0 |  | CVUAS 35069 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.651 | 2.369 | 1 | 0 | 0
| 14 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35069_A8.0 |  | CVUAS 35069 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.645 | 2.421 | 1 | 0 | 0
| 14 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35069_A9.0 |  | CVUAS 35069 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.784 | 2.487 | 1 | 0 | 0
# | 15 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 6528_3_A4.0 |  | CVUAS 6528.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.972 | 2.489 | 1 | 0 | 0
| 15 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 6528_3_A5.0 |  | CVUAS 6528.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.951 | 2.752 | 1 | 0 | 0
| 15 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 6528_3_A6.0 |  | CVUAS 6528.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.978 | 2.671 | 1 | 0 | 0
| 15 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 6528_3_A4.0 |  | CVUAS 6528.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.631 | 2.7 | 1 | 0 | 0
| 15 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 6528_3_A5.0 |  | CVUAS 6528.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.654 | 3.114 | 1 | 0 | 0
| 15 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 6528_3_A6.0 |  | CVUAS 6528.3 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.472 | 2.812 | 1 | 0 | 0
# | 16 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 796_4_B7.0 |  | CVUAS 796.4 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.856 | 2.475 | 1 | 0 | 0
| 16 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 796_4_B8.0 |  | CVUAS 796.4 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.811 | 2.708 | 1 | 0 | 0
| 16 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 796_4_B9.0 |  | CVUAS 796.4 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.694 | 2.88 | 1 | 0 | 0
| 16 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 796_4_B7.0 |  | CVUAS 796.4 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.844 | 2.286 | 1 | 0 | 0
| 16 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 796_4_B8.0 |  | CVUAS 796.4 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.832 | 2.306 | 1 | 0 | 0
| 16 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 796_4_B9.0 |  | CVUAS 796.4 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.822 | 2.371 | 1 | 0 | 0
# | 17 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 10598_E4.0 |  | CVUAS 10598 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.932 | 2.54 | 1 | 0 | 0
| 17 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 10598_E5.0 |  | CVUAS 10598 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.956 | 2.516 | 1 | 0 | 0
| 17 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 10598_E6.0 |  | CVUAS 10598 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.94 | 2.705 | 1 | 0 | 0
| 17 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 10598_E4.0 |  | CVUAS 10598 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.951 | 2.108 | 1 | 0 | 0
| 17 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 10598_E5.0 |  | CVUAS 10598 | FALSE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.944 | 2.433 | 1 | 0 | 0
| 17 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 10598_E6.0 |  | CVUAS 10598 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.937 | 2.529 | 1 | 0 | 0
# | 18 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 10593_D4.0 |  | CVUAS 10593 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.695 | 2.527 | 1 | 0 | 0
| 18 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 10593_D5.0 |  | CVUAS 10593 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.59 | 2.551 | 1 | 0 | 0
| 18 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 10593_D6.0 |  | CVUAS 10593 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.617 | 2.541 | 1 | 0 | 0
| 18 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 10593_D4.0 |  | CVUAS 10593 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.347 | 2.165 | 1 | 0 | 0
| 18 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 10593_D5.0 |  | CVUAS 10593 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.463 | 2.426 | 1 | 0 | 0
| 18 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 10593_D6.0 |  | CVUAS 10593 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.56 | 2.171 | 1 | 0 | 0
# | 19 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34935_E10.0 |  | CVUAS 34935 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.298 | 3.296 | 1 | 0 | 0
| 19 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34935_E11.0 |  | CVUAS 34935 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.201 | 3.715 | 0 | 1 | 0
| 19 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34935_E12.0 |  | CVUAS 34935 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.2 | 3.842 | 0 | 1 | 0
| 19 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34935_E10.0 |  | CVUAS 34935 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.578 | 2.652 | 1 | 0 | 0
| 19 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34935_E11.0 |  | CVUAS 34935 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.29 | 3.403 | 1 | 0 | 0
| 19 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34935_E12.0 |  | CVUAS 34935 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.403 | 3.15 | 1 | 0 | 0
# | 20 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35075_D7.0 |  | CVUAS 35075 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.994 | 2.825 | 1 | 0 | 0
| 20 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35075_D8.0 |  | CVUAS 35075 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.989 | 2.836 | 1 | 0 | 0
| 20 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35075_D9.0 |  | CVUAS 35075 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.984 | 2.749 | 1 | 0 | 0
| 20 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35075_D7.0 |  | CVUAS 35075 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.916 | 2.422 | 1 | 0 | 0
| 20 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35075_D8.0 |  | CVUAS 35075 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.916 | 2.327 | 1 | 0 | 0
| 20 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35075_D9.0 |  | CVUAS 35075 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.924 | 2.276 | 1 | 0 | 0
# | 21 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7232_H4.0 |  | CVUAS 7232 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.874 | 3.135 | 1 | 0 | 0
| 21 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7232_H5.0 |  | CVUAS 7232 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.9 | 3.116 | 1 | 0 | 0
| 21 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7232_H6.0 |  | CVUAS 7232 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.897 | 3.017 | 1 | 0 | 0
| 21 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7232_H4.0 |  | CVUAS 7232 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.89 | 2.51 | 1 | 0 | 0
| 21 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7232_H5.0 |  | CVUAS 7232 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.924 | 2.407 | 1 | 0 | 0
| 21 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7232_H6.0 |  | CVUAS 7232 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.932 | 2.463 | 1 | 0 | 0
# | 22 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32033_D7.0 |  | CVUAS 32033 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.977 | 2.534 | 1 | 0 | 0
| 22 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32033_D8.0 |  | CVUAS 32033 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.979 | 2.68 | 1 | 0 | 0
| 22 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32033_D9.0 |  | CVUAS 32033 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.974 | 2.802 | 1 | 0 | 0
| 22 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32033_D7.0 |  | CVUAS 32033 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.961 | 2.796 | 1 | 0 | 0
| 22 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32033_D8.0 |  | CVUAS 32033 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.979 | 2.58 | 1 | 0 | 0
| 22 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32033_D9.0 |  | CVUAS 32033 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.985 | 2.495 | 1 | 0 | 0
# | 23 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35033_F4.0 |  | CVUAS 35033 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.961 | 2.613 | 1 | 0 | 0
| 23 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35033_F5.0 |  | CVUAS 35033 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.974 | 2.493 | 1 | 0 | 0
| 23 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35033_F6.0 |  | CVUAS 35033 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.961 | 2.695 | 1 | 0 | 0
| 23 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35033_F4.0 |  | CVUAS 35033 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.91 | 2.359 | 1 | 0 | 0
| 23 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35033_F5.0 |  | CVUAS 35033 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.89 | 2.25 | 1 | 0 | 0
| 23 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35033_F6.0 |  | CVUAS 35033 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.125 | 2.032 | 1 | 0 | 0
# | 24 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33538_A7.0 |  | CVUAS 33538 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.906 | 2.658 | 1 | 0 | 0
| 24 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33538_A8.0 |  | CVUAS 33538 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.887 | 2.956 | 1 | 0 | 0
| 24 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33538_A9.0 |  | CVUAS 33538 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.811 | 3.053 | 1 | 0 | 0
| 24 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33538_A7.0 |  | CVUAS 33538 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.785 | 2.109 | 1 | 0 | 0
| 24 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33538_A8.0 |  | CVUAS 33538 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.602 | 2.201 | 1 | 0 | 0
| 24 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33538_A9.0 |  | CVUAS 33538 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.853 | 2.522 | 1 | 0 | 0
# | 25 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35226_B10.0 |  | CVUAS 35226 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.969 | 2.613 | 1 | 0 | 0
| 25 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35226_B11.0 |  | CVUAS 35226 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.97 | 2.769 | 1 | 0 | 0
| 25 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35226_B12.0 |  | CVUAS 35226 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.969 | 2.623 | 1 | 0 | 0
| 25 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35226_B10.0 |  | CVUAS 35226 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.932 | 2.457 | 1 | 0 | 0
| 25 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35226_B11.0 |  | CVUAS 35226 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.913 | 2.46 | 1 | 0 | 0
| 25 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35226_B12.0 |  | CVUAS 35226 | TRUE | O:7 (C1-C4) | O:7 (C1-C4) | 1 | 0.902 | 2.931 | 1 | 0 | 0

## IR BT non-target parm
Method |  | Salmonella O-group differentiation by Bruker salmonella o-groups classifier v3 | [7 empty] | target parameter | O:7 (C1-C4)
Differentation |  | External validation of the a.m. classifier; incubation on COL-SB 37°C 24h | isolates | 142 | [5 empty] | non-target parameter | non - O:7 (C1-C4)
Date of validation |  | see report
|  |  | spectra | 889
[5 empty] | classified | number | percentage |  |  | outlier score | 889
[5 empty] | correct | 854 | 96.06% |  |  | Number in green | 560
[5 empty] | questionable | 35 | 3.937% |  |  | Number in yellow | 294
[5 empty] | wrong | 0 | 0% |  |  | Number in red | 35
[11 empty] | 854 | 35 | 0
Isolate |  | Name (incl. spot no.) | Isolate description | Isolate ID | QT result: wahr=true=pass; falsch=false=fail | Factual classification | IR BT classification | Classifi-cation correkt? (i.e. <> target parm) => 1 | irrelevant! classification score: if <0.6 => invalid | outlier score: <2,5: green, 2,5 - 3,5: yellow, >3,5: red | if outlier score </=3,5 AND classification = factual classification => 1 (correct) | if outlier score >3,5) [classification irrelevant] => 1 (questionable) | if outlier score </=3,5 AND classification <> factual classification => 1 (wrong)
# | 1 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4664_D4.0 |  | CVUAS 4664 | TRUE | non - O:7 (C1-C4) | O:51 | 1 | 0.162 | 2.756 | 1 | 0 | 0
| 1 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4664_D5.0 |  | CVUAS 4664 | TRUE | non - O:7 (C1-C4) | O:51 | 1 | 0.155 | 2.918 | 1 | 0 | 0
| 1 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4664_D6.0 |  | CVUAS 4664 | TRUE | non - O:7 (C1-C4) | O:51 | 1 | 0.149 | 2.761 | 1 | 0 | 0
| 1 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4664_D4.0 |  | CVUAS 4664 | TRUE | non - O:7 (C1-C4) | O:51 | 1 | 0.243 | 2.031 | 1 | 0 | 0
| 1 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4664_D5.0 |  | CVUAS 4664 | TRUE | non - O:7 (C1-C4) | O:51 | 1 | 0.248 | 2.054 | 1 | 0 | 0
| 1 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4664_D6.0 |  | CVUAS 4664 | TRUE | non - O:7 (C1-C4) | O:51 | 1 | 0.196 | 2.146 | 1 | 0 | 0
# | 2 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9887_H10.0 |  | CVUAS 9887 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.413 | 2.524 | 1 | 0 | 0
| 2 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9887_H11.0 |  | CVUAS 9887 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.433 | 2.554 | 1 | 0 | 0
| 2 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9887_H12.0 |  | CVUAS 9887 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.499 | 2.515 | 1 | 0 | 0
| 2 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9887_H10.0 |  | CVUAS 9887 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.467 | 2.528 | 1 | 0 | 0
| 2 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9887_H11.0 |  | CVUAS 9887 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.459 | 2.652 | 1 | 0 | 0
| 2 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9887_H12.0 |  | CVUAS 9887 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.397 | 2.437 | 1 | 0 | 0
# | 3 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 6261_H10.0 |  | CVUAS 6261 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.141 | 2.443 | 1 | 0 | 0
| 3 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 6261_H11.0 |  | CVUAS 6261 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.146 | 2.747 | 1 | 0 | 0
| 3 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 6261_H12.0 |  | CVUAS 6261 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.143 | 2.861 | 1 | 0 | 0
| 3 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 6261_H10.0 |  | CVUAS 6261 | TRUE | non - O:7 (C1-C4) | O:53 | 1 | 0.264 | 1.872 | 1 | 0 | 0
| 3 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 6261_H11.0 |  | CVUAS 6261 | TRUE | non - O:7 (C1-C4) | O:53 | 1 | 0.278 | 1.946 | 1 | 0 | 0
| 3 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 6261_H12.0 |  | CVUAS 6261 | TRUE | non - O:7 (C1-C4) | O:53 | 1 | 0.205 | 1.823 | 1 | 0 | 0
# | 4 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 30491_F10.0 |  | CVUAS 30491 | TRUE | non - O:7 (C1-C4) | O:53 | 1 | 0.515 | 2.128 | 1 | 0 | 0
| 4 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 30491_F11.0 |  | CVUAS 30491 | TRUE | non - O:7 (C1-C4) | O:53 | 1 | 0.368 | 2.032 | 1 | 0 | 0
| 4 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 30491_F12.0 |  | CVUAS 30491 | TRUE | non - O:7 (C1-C4) | O:53 | 1 | 0.468 | 2.066 | 1 | 0 | 0
| 4 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 30491_F10.0 |  | CVUAS 30491 | TRUE | non - O:7 (C1-C4) | O:53 | 1 | 0.733 | 1.788 | 1 | 0 | 0
| 4 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 30491_F11.0 |  | CVUAS 30491 | TRUE | non - O:7 (C1-C4) | O:53 | 1 | 0.426 | 1.738 | 1 | 0 | 0
| 4 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 30491_F12.0 |  | CVUAS 30491 | TRUE | non - O:7 (C1-C4) | O:53 | 1 | 0.607 | 1.815 | 1 | 0 | 0
# | 5 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 3418 _F10.0 |  | CVUAS 3418 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.574 | 2.864 | 1 | 0 | 0
| 5 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 3418 _F11.0 |  | CVUAS 3418 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.163 | 2.977 | 1 | 0 | 0
| 5 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 3418 _F12.0 |  | CVUAS 3418 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.172 | 2.888 | 1 | 0 | 0
| 5 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 3418 _F10.0 |  | CVUAS 3418 | TRUE | non - O:7 (C1-C4) | O:51 | 1 | 0.185 | 3.107 | 1 | 0 | 0
| 5 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 3418 _F11.0 |  | CVUAS 3418 | TRUE | non - O:7 (C1-C4) | O:51 | 1 | 0.18 | 2.981 | 1 | 0 | 0
| 5 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 3418 _F12.0 |  | CVUAS 3418 | TRUE | non - O:7 (C1-C4) | O:51 | 1 | 0.175 | 3.102 | 1 | 0 | 0
# | 6 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 6615_B4.0 |  | CVUAS 6615 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.585 | 1.751 | 1 | 0 | 0
| 6 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 6615_B5.0 |  | CVUAS 6615 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.504 | 1.969 | 1 | 0 | 0
| 6 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 6615_B6.0 |  | CVUAS 6615 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.495 | 2.171 | 1 | 0 | 0
| 6 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 6615_B4.0 |  | CVUAS 6615 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.856 | 1.724 | 1 | 0 | 0
| 6 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 6615_B5.0 |  | CVUAS 6615 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.795 | 1.97 | 1 | 0 | 0
| 6 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 6615_B6.0 |  | CVUAS 6615 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.751 | 1.747 | 1 | 0 | 0
# | 7 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 8523_A10.0 |  | CVUAS 8523 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.275 | 2.466 | 1 | 0 | 0
| 7 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 8523_A11.0 |  | CVUAS 8523 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.317 | 2.429 | 1 | 0 | 0
| 7 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 8523_A12.0 |  | CVUAS 8523 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.332 | 2.352 | 1 | 0 | 0
| 7 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 8523_A10.0 |  | CVUAS 8523 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.687 | 1.984 | 1 | 0 | 0
| 7 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 8523_A11.0 |  | CVUAS 8523 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.542 | 2.194 | 1 | 0 | 0
| 7 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 8523_A12.0 |  | CVUAS 8523 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.482 | 2.412 | 1 | 0 | 0
# | 8 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32739_B10.0 |  | CVUAS 32739 | TRUE | non - O:7 (C1-C4) | O:57 | 1 | 0.114 | 2.07 | 1 | 0 | 0
| 8 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32739_B11.0 |  | CVUAS 32739 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.133 | 2.287 | 1 | 0 | 0
| 8 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32739_B12.0 |  | CVUAS 32739 | TRUE | non - O:7 (C1-C4) | O:42 (T) | 1 | 0.1 | 2.394 | 1 | 0 | 0
| 8 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32739_B10.0 |  | CVUAS 32739 | TRUE | non - O:7 (C1-C4) | O:57 | 1 | 0.225 | 2.446 | 1 | 0 | 0
| 8 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32739_B11.0 |  | CVUAS 32739 | TRUE | non - O:7 (C1-C4) | O:57 | 1 | 0.215 | 2.57 | 1 | 0 | 0
| 8 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32739_B12.0 |  | CVUAS 32739 | TRUE | non - O:7 (C1-C4) | O:57 | 1 | 0.213 | 2.544 | 1 | 0 | 0
# | 9 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 884_2_C7.0 |  | CVUAS 884.2 | TRUE | non - O:7 (C1-C4) | O:57 | 1 | 0.461 | 1.797 | 1 | 0 | 0
| 9 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 884_2_C8.0 |  | CVUAS 884.2 | TRUE | non - O:7 (C1-C4) | O:57 | 1 | 0.369 | 1.922 | 1 | 0 | 0
| 9 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 884_2_C9.0 |  | CVUAS 884.2 | TRUE | non - O:7 (C1-C4) | O:57 | 1 | 0.315 | 1.936 | 1 | 0 | 0
| 9 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 884_2_C7.0 |  | CVUAS 884.2 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.351 | 2.028 | 1 | 0 | 0
| 9 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 884_2_C8.0 |  | CVUAS 884.2 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.261 | 2.116 | 1 | 0 | 0
| 9 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 884_2_C9.0 |  | CVUAS 884.2 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.228 | 2.116 | 1 | 0 | 0
# | 10 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35223_A10.0 |  | CVUAS 35223 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.131 | 3.131 | 1 | 0 | 0
| 10 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35223_A11.0 |  | CVUAS 35223 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.14 | 2.986 | 1 | 0 | 0
| 10 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35223_A12.0 |  | CVUAS 35223 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.173 | 3.192 | 1 | 0 | 0
| 10 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35223_A10.0 |  | CVUAS 35223 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.35 | 3.661 | 0 | 1 | 0
| 10 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35223_A11.0 |  | CVUAS 35223 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.235 | 3.505 | 0 | 1 | 0
| 10 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35223_A12.0 |  | CVUAS 35223 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.365 | 3.724 | 0 | 1 | 0
# | 11 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35403_C10.0 |  | CVUAS 35403 | TRUE | non - O:7 (C1-C4) | O:58 | 1 | 0.626 | 3.061 | 1 | 0 | 0
| 11 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35403_C11.0 |  | CVUAS 35403 | TRUE | non - O:7 (C1-C4) | O:58 | 1 | 0.649 | 2.79 | 1 | 0 | 0
| 11 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35403_C12.0 |  | CVUAS 35403 | TRUE | non - O:7 (C1-C4) | O:58 | 1 | 0.618 | 2.874 | 1 | 0 | 0
| 11 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35403_C10.0 |  | CVUAS 35403 | TRUE | non - O:7 (C1-C4) | O:58 | 1 | 0.613 | 2.318 | 1 | 0 | 0
| 11 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35403_C11.0 |  | CVUAS 35403 | TRUE | non - O:7 (C1-C4) | O:58 | 1 | 0.55 | 2.244 | 1 | 0 | 0
| 11 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35403_C12.0 |  | CVUAS 35403 | TRUE | non - O:7 (C1-C4) | O:58 | 1 | 0.6 | 2.299 | 1 | 0 | 0
# | 12 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5598_2_D10.0 |  | CVUAS 5598.2 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.24 | 3.412 | 1 | 0 | 0
| 12 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5598_2_D11.0 |  | CVUAS 5598.2 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.246 | 3.535 | 0 | 1 | 0
| 12 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5598_2_D12.0 |  | CVUAS 5598.2 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.241 | 3.505 | 0 | 1 | 0
| 12 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5598_2_D10.0 |  | CVUAS 5598.2 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.254 | 3.701 | 0 | 1 | 0
| 12 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5598_2_D11.0 |  | CVUAS 5598.2 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.247 | 3.574 | 0 | 1 | 0
| 12 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5598_2_D12.0 |  | CVUAS 5598.2 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.283 | 3.719 | 0 | 1 | 0
# | 13 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 8304_2_G7.0 |  | CVUAS 8304.2 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.278 | 3.865 | 0 | 1 | 0
| 13 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 8304_2_G8.0 |  | CVUAS 8304.2 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.239 | 3.93 | 0 | 1 | 0
| 13 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 8304_2_G9.0 |  | CVUAS 8304.2 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.224 | 3.914 | 0 | 1 | 0
| 13 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 8304_2_G7.0 |  | CVUAS 8304.2 | FALSE | non - O:7 (C1-C4) | O:60 | 1 | 0.319 | 4.009 | 0 | 1 | 0
| 13 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 8304_2_G8.0 |  | CVUAS 8304.2 | FALSE | non - O:7 (C1-C4) | O:60 | 1 | 0.611 | 4.531 | 0 | 1 | 0
| 13 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 8304_2_G9.0 |  | CVUAS 8304.2 | FALSE | non - O:7 (C1-C4) | O:60 | 1 | 0.495 | 4.203 | 0 | 1 | 0
| 13 | 231221-1432-BayP-Sal Vali WDH_CVUAS 8304_2_E4.0 |  | CVUAS 8304.2 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.331 | 3.524 | 0 | 1 | 0
| 13 | 231221-1432-BayP-Sal Vali WDH_CVUAS 8304_2_E5.0 |  | CVUAS 8304.2 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.233 | 3.863 | 0 | 1 | 0
| 13 | 231221-1432-BayP-Sal Vali WDH_CVUAS 8304_2_E6.0 |  | CVUAS 8304.2 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.255 | 3.9 | 0 | 1 | 0
# | 14 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34995_H10.0 |  | CVUAS 34995 | TRUE | non - O:7 (C1-C4) | O:16 (I) | 1 | 0.321 | 3.976 | 0 | 1 | 0
| 14 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34995_H11.0 |  | CVUAS 34995 | TRUE | non - O:7 (C1-C4) | O:35 (O) | 1 | 0.148 | 2.406 | 1 | 0 | 0
| 14 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34995_H12.0 |  | CVUAS 34995 | TRUE | non - O:7 (C1-C4) | O:60 | 1 | 0.21 | 2.369 | 1 | 0 | 0
| 14 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34995_H10.0 |  | CVUAS 34995 | TRUE | non - O:7 (C1-C4) | O:60 | 1 | 0.491 | 2.178 | 1 | 0 | 0
| 14 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34995_H11.0 |  | CVUAS 34995 | TRUE | non - O:7 (C1-C4) | O:60 | 1 | 0.275 | 2.217 | 1 | 0 | 0
| 14 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34995_H12.0 |  | CVUAS 34995 | TRUE | non - O:7 (C1-C4) | O:60 | 1 | 0.473 | 2.272 | 1 | 0 | 0
# | 15 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35583_F10.0 |  | CVUAS 35583 | TRUE | non - O:7 (C1-C4) | O:60 | 1 | 0.479 | 2.507 | 1 | 0 | 0
| 15 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35583_F11.0 |  | CVUAS 35583 | TRUE | non - O:7 (C1-C4) | O:60 | 1 | 0.412 | 2.704 | 1 | 0 | 0
| 15 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35583_F12.0 |  | CVUAS 35583 | TRUE | non - O:7 (C1-C4) | O:60 | 1 | 0.346 | 2.717 | 1 | 0 | 0
| 15 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35583_F10.0 |  | CVUAS 35583 | TRUE | non - O:7 (C1-C4) | O:60 | 1 | 0.534 | 2.347 | 1 | 0 | 0
| 15 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35583_F11.0 |  | CVUAS 35583 | TRUE | non - O:7 (C1-C4) | O:60 | 1 | 0.428 | 2.648 | 1 | 0 | 0
| 15 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35583_F12.0 |  | CVUAS 35583 | TRUE | non - O:7 (C1-C4) | O:60 | 1 | 0.455 | 2.499 | 1 | 0 | 0
# | 16 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35494_E10.0 |  | CVUAS 35494 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.122 | 3.31 | 1 | 0 | 0
| 16 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35494_E11.0 |  | CVUAS 35494 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.131 | 3.319 | 1 | 0 | 0
| 16 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35494_E12.0 |  | CVUAS 35494 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.133 | 3.266 | 1 | 0 | 0
| 16 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35494_E10.0 |  | CVUAS 35494 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.143 | 3.02 | 1 | 0 | 0
| 16 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35494_E11.0 |  | CVUAS 35494 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.302 | 3.247 | 1 | 0 | 0
| 16 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35494_E12.0 |  | CVUAS 35494 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.143 | 2.99 | 1 | 0 | 0
# | 17 | 231212-1451-Loel Sal Vali Gr 7_ CVUAS 35067_G4.0 |  | CVUAS 35067 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.171 | 3.386 | 1 | 0 | 0
| 17 | 231212-1451-Loel Sal Vali Gr 7_ CVUAS 35067_G5.0 |  | CVUAS 35067 | TRUE | non - O:7 (C1-C4) | O:61 | 1 | 0.183 | 3.393 | 1 | 0 | 0
| 17 | 231212-1451-Loel Sal Vali Gr 7_ CVUAS 35067_G6.0 |  | CVUAS 35067 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.167 | 3.538 | 0 | 1 | 0
| 17 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35067_G4.0 |  | CVUAS 35067 | TRUE | non - O:7 (C1-C4) | O:61 | 1 | 0.188 | 2.676 | 1 | 0 | 0
| 17 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35067_G5.0 |  | CVUAS 35067 | TRUE | non - O:7 (C1-C4) | O:61 | 1 | 0.227 | 2.923 | 1 | 0 | 0
| 17 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35067_G6.0 |  | CVUAS 35067 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.144 | 2.711 | 1 | 0 | 0
# | 18 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 452_2 _F4.0 |  | CVUAS 452.2 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.824 | 1.954 | 1 | 0 | 0
| 18 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 452_2 _F5.0 |  | CVUAS 452.2 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.763 | 1.926 | 1 | 0 | 0
| 18 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 452_2 _F6.0 |  | CVUAS 452.2 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.637 | 2.083 | 1 | 0 | 0
| 18 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 452_2 _F4.0 |  | CVUAS 452.2 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.617 | 2.025 | 1 | 0 | 0
| 18 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 452_2 _F5.0 |  | CVUAS 452.2 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.489 | 2.18 | 1 | 0 | 0
| 18 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 452_2 _F6.0 |  | CVUAS 452.2 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.645 | 1.925 | 1 | 0 | 0
# | 19 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 9888_A4.0 |  | CVUAS 9888 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.714 | 1.681 | 1 | 0 | 0
| 19 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 9888_A5.0 |  | CVUAS 9888 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.699 | 1.89 | 1 | 0 | 0
| 19 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 9888_A6.0 |  | CVUAS 9888 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.737 | 1.818 | 1 | 0 | 0
| 19 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 9888_A4.0 |  | CVUAS 9888 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.696 | 2.077 | 1 | 0 | 0
| 19 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 9888_A5.0 |  | CVUAS 9888 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.665 | 2.14 | 1 | 0 | 0
| 19 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 9888_A6.0 |  | CVUAS 9888 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.678 | 2.176 | 1 | 0 | 0
# | 20 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35493_D10.0 |  | CVUAS 35493 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.997 | 2.585 | 1 | 0 | 0
| 20 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35493_D11.0 |  | CVUAS 35493 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.996 | 2.416 | 1 | 0 | 0
| 20 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35493_D12.0 |  | CVUAS 35493 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.993 | 2.748 | 1 | 0 | 0
| 20 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35493_D10.0 |  | CVUAS 35493 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.995 | 2.332 | 1 | 0 | 0
| 20 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35493_D11.0 |  | CVUAS 35493 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.997 | 2.42 | 1 | 0 | 0
| 20 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35493_D12.0 |  | CVUAS 35493 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.997 | 2.451 | 1 | 0 | 0
# | 21 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7666_2_B7.0 |  | CVUAS 7666.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.998 | 3.073 | 1 | 0 | 0
| 21 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7666_2_B8.0 |  | CVUAS 7666.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.989 | 2.84 | 1 | 0 | 0
| 21 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7666_2_B9.0 |  | CVUAS 7666.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.967 | 2.62 | 1 | 0 | 0
| 21 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7666_2_B7.0 |  | CVUAS 7666.2 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.953 | 2.939 | 1 | 0 | 0
| 21 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7666_2_B8.0 |  | CVUAS 7666.2 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.981 | 3.146 | 1 | 0 | 0
| 21 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7666_2_B9.0 |  | CVUAS 7666.2 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.959 | 3.145 | 1 | 0 | 0
| 21 | 231221-1432-BayP-Sal Vali WDH_CVUAS 7666_2_B4.0 |  | CVUAS 7666.2 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.51 | 1 | 0 | 0
| 21 | 231221-1432-BayP-Sal Vali WDH_CVUAS 7666_2_B5.0 |  | CVUAS 7666.2 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.588 | 1 | 0 | 0
| 21 | 231221-1432-BayP-Sal Vali WDH_CVUAS 7666_2_B6.0 |  | CVUAS 7666.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.727 | 1 | 0 | 0
# | 22 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 1075_4_E7.0 |  | CVUAS 1075.4 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.996 | 1.884 | 1 | 0 | 0
| 22 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 1075_4_E8.0 |  | CVUAS 1075.4 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.994 | 1.628 | 1 | 0 | 0
| 22 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 1075_4_E9.0 |  | CVUAS 1075.4 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.994 | 1.759 | 1 | 0 | 0
| 22 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 1075_4_E7.0 |  | CVUAS 1075.4 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 1 | 2.864 | 1 | 0 | 0
| 22 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 1075_4_E8.0 |  | CVUAS 1075.4 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 1 | 2.141 | 1 | 0 | 0
| 22 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 1075_4_E9.0 |  | CVUAS 1075.4 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 1 | 2.284 | 1 | 0 | 0
# | 23 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5870_F10.0 |  | CVUAS 5870 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.99 | 2.766 | 1 | 0 | 0
| 23 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5870_F11.0 |  | CVUAS 5870 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.994 | 2.699 | 1 | 0 | 0
| 23 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5870_F12.0 |  | CVUAS 5870 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.995 | 2.61 | 1 | 0 | 0
| 23 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5870_F10.0 |  | CVUAS 5870 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.976 | 1.93 | 1 | 0 | 0
| 23 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5870_F11.0 |  | CVUAS 5870 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.987 | 1.73 | 1 | 0 | 0
| 23 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5870_F12.0 |  | CVUAS 5870 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.981 | 1.952 | 1 | 0 | 0
# | 24 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31295_2_G4.0 |  | CVUAS 31295.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.716 | 1 | 0 | 0
| 24 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31295_2_G5.0 |  | CVUAS 31295.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 1 | 2.825 | 1 | 0 | 0
| 24 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31295_2_G6.0 |  | CVUAS 31295.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.802 | 1 | 0 | 0
| 24 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31295_2_G4.0 |  | CVUAS 31295.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.985 | 2.427 | 1 | 0 | 0
| 24 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31295_2_G5.0 |  | CVUAS 31295.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.985 | 2.238 | 1 | 0 | 0
| 24 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31295_2_G6.0 |  | CVUAS 31295.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.982 | 2.388 | 1 | 0 | 0
# | 25 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32369_F7.0 |  | CVUAS 32369 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 1 | 2.522 | 1 | 0 | 0
| 25 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32369_F8.0 |  | CVUAS 32369 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.993 | 2.354 | 1 | 0 | 0
| 25 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32369_F9.0 |  | CVUAS 32369 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.996 | 2.42 | 1 | 0 | 0
| 25 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32369_F7.0 |  | CVUAS 32369 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 1 | 2.35 | 1 | 0 | 0
| 25 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32369_F8.0 |  | CVUAS 32369 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 1 | 2.54 | 1 | 0 | 0
| 25 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32369_F9.0 |  | CVUAS 32369 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 1 | 2.379 | 1 | 0 | 0
# | 26 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 30692_B4.0 |  | CVUAS 30692 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.953 | 2.717 | 1 | 0 | 0
| 26 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 30692_B5.0 |  | CVUAS 30692 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.93 | 2.606 | 1 | 0 | 0
| 26 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 30692_B6.0 |  | CVUAS 30692 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.993 | 2.487 | 1 | 0 | 0
| 26 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 30692_B4.0 |  | CVUAS 30692 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.966 | 2.208 | 1 | 0 | 0
| 26 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 30692_B5.0 |  | CVUAS 30692 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.965 | 2.335 | 1 | 0 | 0
| 26 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 30692_B6.0 |  | CVUAS 30692 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.975 | 2.268 | 1 | 0 | 0
# | 27 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34975_G10.0 |  | CVUAS 34975 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.888 | 1 | 0 | 0
| 27 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34975_G11.0 |  | CVUAS 34975 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.833 | 1 | 0 | 0
| 27 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34975_G12.0 |  | CVUAS 34975 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.473 | 1 | 0 | 0
| 27 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34975_G10.0 |  | CVUAS 34975 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.751 | 1 | 0 | 0
| 27 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34975_G11.0 |  | CVUAS 34975 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.573 | 1 | 0 | 0
| 27 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34975_G12.0 |  | CVUAS 34975 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.436 | 1 | 0 | 0
# | 28 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 34997 _B4.0 |  | CVUAS 34997 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.989 | 2.319 | 1 | 0 | 0
| 28 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 34997 _B5.0 |  | CVUAS 34997 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.994 | 2.455 | 1 | 0 | 0
| 28 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 34997 _B6.0 |  | CVUAS 34997 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.997 | 2.534 | 1 | 0 | 0
| 28 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 34997 _B4.0 |  | CVUAS 34997 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.981 | 1.916 | 1 | 0 | 0
| 28 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 34997 _B5.0 |  | CVUAS 34997 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.99 | 2.332 | 1 | 0 | 0
| 28 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 34997 _B6.0 |  | CVUAS 34997 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.983 | 2.413 | 1 | 0 | 0
# | 29 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11719_C10.0 |  | CVUAS 11719 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.779 | 1 | 0 | 0
| 29 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11719_C11.0 |  | CVUAS 11719 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.669 | 1 | 0 | 0
| 29 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11719_C12.0 |  | CVUAS 11719 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.998 | 2.493 | 1 | 0 | 0
| 29 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11719_C10.0 |  | CVUAS 11719 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.025 | 1 | 0 | 0
| 29 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11719_C11.0 |  | CVUAS 11719 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 1.921 | 1 | 0 | 0
| 29 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11719_C12.0 |  | CVUAS 11719 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.413 | 1 | 0 | 0
| 29 | 231221-1432-BayP-Sal Vali WDH_CVUAS 11719_B7.0 |  | CVUAS 11719 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.995 | 2.425 | 1 | 0 | 0
| 29 | 231221-1432-BayP-Sal Vali WDH_CVUAS 11719_B8.0 |  | CVUAS 11719 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.996 | 2.236 | 1 | 0 | 0
| 29 | 231221-1432-BayP-Sal Vali WDH_CVUAS 11719_B9.0 |  | CVUAS 11719 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.994 | 2.477 | 1 | 0 | 0
# | 30 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33884_D7.0 |  | CVUAS 33884 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.998 | 1.493 | 1 | 0 | 0
| 30 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33884_D8.0 |  | CVUAS 33884 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.996 | 1.49 | 1 | 0 | 0
| 30 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33884_D9.0 |  | CVUAS 33884 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.996 | 1.512 | 1 | 0 | 0
| 30 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33884_D7.0 |  | CVUAS 33884 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.995 | 1.705 | 1 | 0 | 0
| 30 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33884_D8.0 |  | CVUAS 33884 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.997 | 1.842 | 1 | 0 | 0
| 30 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33884_D9.0 |  | CVUAS 33884 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.995 | 1.657 | 1 | 0 | 0
# | 31 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33881_C7.0 |  | CVUAS 33881 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.997 | 2.413 | 1 | 0 | 0
| 31 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33881_C8.0 |  | CVUAS 33881 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 1 | 2.775 | 1 | 0 | 0
| 31 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33881_C9.0 |  | CVUAS 33881 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.995 | 2.448 | 1 | 0 | 0
| 31 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33881_C7.0 |  | CVUAS 33881 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.97 | 2.344 | 1 | 0 | 0
| 31 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33881_C8.0 |  | CVUAS 33881 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.974 | 2.379 | 1 | 0 | 0
| 31 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33881_C9.0 |  | CVUAS 33881 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.981 | 2.369 | 1 | 0 | 0
# | 32 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32051_E7.0 |  | CVUAS 32051 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 1 | 2.324 | 1 | 0 | 0
| 32 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32051_E8.0 |  | CVUAS 32051 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 1 | 2.425 | 1 | 0 | 0
| 32 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32051_E9.0 |  | CVUAS 32051 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 1 | 2.363 | 1 | 0 | 0
| 32 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32051_E7.0 |  | CVUAS 32051 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.989 | 1.959 | 1 | 0 | 0
| 32 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32051_E8.0 |  | CVUAS 32051 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.99 | 1.739 | 1 | 0 | 0
| 32 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32051_E9.0 |  | CVUAS 32051 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.994 | 1.766 | 1 | 0 | 0
# | 33 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32797_D10.0 |  | CVUAS 32797 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.997 | 2.338 | 1 | 0 | 0
| 33 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32797_D11.0 |  | CVUAS 32797 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.996 | 2.565 | 1 | 0 | 0
| 33 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32797_D12.0 |  | CVUAS 32797 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.998 | 2.393 | 1 | 0 | 0
| 33 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32797_D10.0 |  | CVUAS 32797 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.998 | 2.325 | 1 | 0 | 0
| 33 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32797_D11.0 |  | CVUAS 32797 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.491 | 1 | 0 | 0
| 33 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32797_D12.0 |  | CVUAS 32797 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.658 | 1 | 0 | 0
# | 34 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 9912_2_B4.0 |  | CVUAS 9912.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.978 | 2.074 | 1 | 0 | 0
| 34 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 9912_2_B5.0 |  | CVUAS 9912.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.986 | 2.038 | 1 | 0 | 0
| 34 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 9912_2_B6.0 |  | CVUAS 9912.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.991 | 2.058 | 1 | 0 | 0
| 34 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 9912_2_B4.0 |  | CVUAS 9912.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.981 | 2.098 | 1 | 0 | 0
| 34 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 9912_2_B5.0 |  | CVUAS 9912.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.974 | 2.207 | 1 | 0 | 0
| 34 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 9912_2_B6.0 |  | CVUAS 9912.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.984 | 2.015 | 1 | 0 | 0
# | 35 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33381_G4.0 |  | CVUAS 33381 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.941 | 2.319 | 1 | 0 | 0
| 35 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33381_G5.0 |  | CVUAS 33381 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.951 | 2.45 | 1 | 0 | 0
| 35 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33381_G6.0 |  | CVUAS 33381 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.985 | 2.4 | 1 | 0 | 0
| 35 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33381_G4.0 |  | CVUAS 33381 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.993 | 2.159 | 1 | 0 | 0
| 35 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33381_G5.0 |  | CVUAS 33381 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.981 | 2.047 | 1 | 0 | 0
| 35 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33381_G6.0 |  | CVUAS 33381 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.992 | 2.207 | 1 | 0 | 0
# | 36 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33385_H4.0 |  | CVUAS 33385 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.952 | 2.032 | 1 | 0 | 0
| 36 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33385_H5.0 |  | CVUAS 33385 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.98 | 2.043 | 1 | 0 | 0
| 36 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33385_H6.0 |  | CVUAS 33385 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.977 | 2.066 | 1 | 0 | 0
| 36 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33385_H4.0 |  | CVUAS 33385 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.985 | 1.784 | 1 | 0 | 0
| 36 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33385_H5.0 |  | CVUAS 33385 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.983 | 1.834 | 1 | 0 | 0
| 36 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33385_H6.0 |  | CVUAS 33385 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.994 | 2.02 | 1 | 0 | 0
# | 37 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34398_G7.0 |  | CVUAS 34398 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 1 | 2.421 | 1 | 0 | 0
| 37 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34398_G8.0 |  | CVUAS 34398 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 1 | 2.464 | 1 | 0 | 0
| 37 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34398_G9.0 |  | CVUAS 34398 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.508 | 1 | 0 | 0
| 37 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34398_G7.0 |  | CVUAS 34398 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.994 | 1.89 | 1 | 0 | 0
| 37 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34398_G8.0 |  | CVUAS 34398 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.998 | 2.691 | 1 | 0 | 0
| 37 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34398_G9.0 |  | CVUAS 34398 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.997 | 2.022 | 1 | 0 | 0
# | 38 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34886_B10.0 |  | CVUAS 34886 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.998 | 1.873 | 1 | 0 | 0
| 38 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34886_B11.0 |  | CVUAS 34886 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.998 | 1.826 | 1 | 0 | 0
| 38 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34886_B12.0 |  | CVUAS 34886 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 1.928 | 1 | 0 | 0
| 38 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34886_B10.0 |  | CVUAS 34886 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.979 | 1.547 | 1 | 0 | 0
| 38 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34886_B11.0 |  | CVUAS 34886 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.975 | 1.63 | 1 | 0 | 0
| 38 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34886_B12.0 |  | CVUAS 34886 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.988 | 1.749 | 1 | 0 | 0
# | 39 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34934_D10.0 |  | CVUAS 34934 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 1.989 | 1 | 0 | 0
| 39 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34934_D11.0 |  | CVUAS 34934 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.209 | 1 | 0 | 0
| 39 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34934_D12.0 |  | CVUAS 34934 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.277 | 1 | 0 | 0
| 39 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34934_D10.0 |  | CVUAS 34934 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.997 | 2.066 | 1 | 0 | 0
| 39 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34934_D11.0 |  | CVUAS 34934 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.998 | 1.801 | 1 | 0 | 0
| 39 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34934_D12.0 |  | CVUAS 34934 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 1.993 | 1 | 0 | 0
# | 40 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35013_E4.0 |  | CVUAS 35013 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.99 | 2.408 | 1 | 0 | 0
| 40 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35013_E5.0 |  | CVUAS 35013 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.998 | 1.797 | 1 | 0 | 0
| 40 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35013_E6.0 |  | CVUAS 35013 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 1.874 | 1 | 0 | 0
| 40 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35013_E4.0 |  | CVUAS 35013 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.972 | 1.868 | 1 | 0 | 0
| 40 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35013_E5.0 |  | CVUAS 35013 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.974 | 1.931 | 1 | 0 | 0
| 40 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35013_E6.0 |  | CVUAS 35013 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.934 | 2.163 | 1 | 0 | 0
# | 41 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 9929_C4.0 |  | CVUAS 9929 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.968 | 2.382 | 1 | 0 | 0
| 41 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 9929_C5.0 |  | CVUAS 9929 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.986 | 2.237 | 1 | 0 | 0
| 41 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 9929_C6.0 |  | CVUAS 9929 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.982 | 2.339 | 1 | 0 | 0
| 41 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 9929_C4.0 |  | CVUAS 9929 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.979 | 1.937 | 1 | 0 | 0
| 41 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 9929_C5.0 |  | CVUAS 9929 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.997 | 2.163 | 1 | 0 | 0
| 41 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 9929_C6.0 |  | CVUAS 9929 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.997 | 2.396 | 1 | 0 | 0
| 41 | 231222-1357-BayP-Sal Vali WDH_CVUAS 9929_A4.0 |  | CVUAS 9929 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.998 | 2.343 | 1 | 0 | 0
| 41 | 231222-1357-BayP-Sal Vali WDH_CVUAS 9929_A5.0 |  | CVUAS 9929 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.997 | 2.261 | 1 | 0 | 0
| 41 | 231222-1357-BayP-Sal Vali WDH_CVUAS 9929_A6.0 |  | CVUAS 9929 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.996 | 2.028 | 1 | 0 | 0
# | 42 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11067_3_G4.0 |  | CVUAS 11067.3 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.907 | 2.418 | 1 | 0 | 0
| 42 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11067_3_G5.0 |  | CVUAS 11067.3 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.987 | 2.063 | 1 | 0 | 0
| 42 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11067_3_G6.0 |  | CVUAS 11067.3 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.997 | 1.992 | 1 | 0 | 0
| 42 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11067_3_G4.0 |  | CVUAS 11067.3 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.958 | 2.008 | 1 | 0 | 0
| 42 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11067_3_G5.0 |  | CVUAS 11067.3 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.998 | 2.687 | 1 | 0 | 0
| 42 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11067_3_G6.0 |  | CVUAS 11067.3 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.572 | 1 | 0 | 0
| 42 | 231222-1357-BayP-Sal Vali WDH_CVUAS 11067_3_B4.0 |  | CVUAS 11067.3 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.216 | 1 | 0 | 0
| 42 | 231222-1357-BayP-Sal Vali WDH_CVUAS 11067_3_B5.0 |  | CVUAS 11067.3 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.995 | 2.26 | 1 | 0 | 0
| 42 | 231222-1357-BayP-Sal Vali WDH_CVUAS 11067_3_B6.0 |  | CVUAS 11067.3 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.992 | 2.054 | 1 | 0 | 0
# | 43 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 35875_G7.0 |  | CVUAS 35875 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 1.867 | 1 | 0 | 0
| 43 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 35875_G8.0 |  | CVUAS 35875 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 1.768 | 1 | 0 | 0
| 43 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 35875_G9.0 |  | CVUAS 35875 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 1.802 | 1 | 0 | 0
| 43 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 35875_G7.0 |  | CVUAS 35875 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.998 | 2.17 | 1 | 0 | 0
| 43 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 35875_G8.0 |  | CVUAS 35875 | FALSE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.998 | 2.335 | 1 | 0 | 0
| 43 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 35875_G9.0 |  | CVUAS 35875 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.995 | 2.211 | 1 | 0 | 0
# | 44 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5601_3_E10.0 |  | CVUAS 5601.3 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 1 | 2.504 | 1 | 0 | 0
| 44 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5601_3_E11.0 |  | CVUAS 5601.3 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.999 | 2.61 | 1 | 0 | 0
| 44 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5601_3_E12.0 |  | CVUAS 5601.3 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 1 | 2.609 | 1 | 0 | 0
| 44 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5601_3_E10.0 |  | CVUAS 5601.3 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.982 | 2.272 | 1 | 0 | 0
| 44 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5601_3_E11.0 |  | CVUAS 5601.3 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.972 | 2.3 | 1 | 0 | 0
| 44 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5601_3_E12.0 |  | CVUAS 5601.3 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.983 | 2.098 | 1 | 0 | 0
# | 45 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4610_C4.0 |  | CVUAS 4610 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.966 | 1.919 | 1 | 0 | 0
| 45 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4610_C5.0 |  | CVUAS 4610 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.96 | 1.996 | 1 | 0 | 0
| 45 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4610_C6.0 |  | CVUAS 4610 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.971 | 2.061 | 1 | 0 | 0
| 45 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4610_C4.0 |  | CVUAS 4610 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.967 | 1.935 | 1 | 0 | 0
| 45 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4610_C5.0 |  | CVUAS 4610 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.982 | 2.013 | 1 | 0 | 0
| 45 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4610_C6.0 |  | CVUAS 4610 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.982 | 2.117 | 1 | 0 | 0
# | 46 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5004_2_D7.0 |  | CVUAS 5004.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.997 | 2.151 | 1 | 0 | 0
| 46 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5004_2_D8.0 |  | CVUAS 5004.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.998 | 2.447 | 1 | 0 | 0
| 46 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5004_2_D9.0 |  | CVUAS 5004.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.998 | 2.547 | 1 | 0 | 0
| 46 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5004_2_D7.0 |  | CVUAS 5004.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.994 | 1.8 | 1 | 0 | 0
| 46 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5004_2_D8.0 |  | CVUAS 5004.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.993 | 1.997 | 1 | 0 | 0
| 46 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5004_2_D9.0 |  | CVUAS 5004.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.994 | 1.921 | 1 | 0 | 0
# | 47 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7940_2_D7.0 |  | CVUAS 7940.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.988 | 2.744 | 1 | 0 | 0
| 47 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7940_2_D8.0 |  | CVUAS 7940.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.988 | 2.893 | 1 | 0 | 0
| 47 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7940_2_D9.0 |  | CVUAS 7940.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.991 | 2.346 | 1 | 0 | 0
| 47 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7940_2_D7.0 |  | CVUAS 7940.2 | FALSE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.995 | 3.431 | 1 | 0 | 0
| 47 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7940_2_D8.0 |  | CVUAS 7940.2 | FALSE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.99 | 2.866 | 1 | 0 | 0
| 47 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7940_2_D9.0 |  | CVUAS 7940.2 | FALSE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.985 | 3.583 | 0 | 1 | 0
| 47 | 231221-1432-BayP-Sal Vali WDH_CVUAS 7940_2_D4.0 |  | CVUAS 7940.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.979 | 2.076 | 1 | 0 | 0
| 47 | 231221-1432-BayP-Sal Vali WDH_CVUAS 7940_2_D5.0 |  | CVUAS 7940.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.974 | 2.619 | 1 | 0 | 0
| 47 | 231221-1432-BayP-Sal Vali WDH_CVUAS 7940_2_D6.0 |  | CVUAS 7940.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.965 | 2.613 | 1 | 0 | 0
# | 48 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9383_D10.0 |  | CVUAS 9383 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.99 | 2.628 | 1 | 0 | 0
| 48 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9383_D11.0 |  | CVUAS 9383 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.99 | 2.588 | 1 | 0 | 0
| 48 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9383_D12.0 |  | CVUAS 9383 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.991 | 2.684 | 1 | 0 | 0
| 48 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9383_D10.0 |  | CVUAS 9383 | FALSE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.99 | 2.531 | 1 | 0 | 0
| 48 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9383_D11.0 |  | CVUAS 9383 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.997 | 2.3 | 1 | 0 | 0
| 48 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9383_D12.0 |  | CVUAS 9383 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.994 | 2.65 | 1 | 0 | 0
# | 49 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 34996_A4.0 |  | CVUAS 34996 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.995 | 2.057 | 1 | 0 | 0
| 49 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 34996_A5.0 |  | CVUAS 34996 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.994 | 2.324 | 1 | 0 | 0
| 49 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 34996_A6.0 |  | CVUAS 34996 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.996 | 2.065 | 1 | 0 | 0
| 49 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 34996_A4.0 |  | CVUAS 34996 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.997 | 1.711 | 1 | 0 | 0
| 49 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 34996_A5.0 |  | CVUAS 34996 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.998 | 1.958 | 1 | 0 | 0
| 49 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 34996_A6.0 |  | CVUAS 34996 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.998 | 1.912 | 1 | 0 | 0
# | 50 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 4586_2_A10.0 |  | CVUAS 4586.2 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.647 | 1.881 | 1 | 0 | 0
| 50 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 4586_2_A11.0 |  | CVUAS 4586.2 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.353 | 2.061 | 1 | 0 | 0
| 50 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 4586_2_A12.0 |  | CVUAS 4586.2 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.144 | 2.218 | 1 | 0 | 0
| 50 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 4586_2_E10.0 |  | CVUAS 4586.2 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.246 | 1.731 | 1 | 0 | 0
| 50 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 4586_2_E11.0 |  | CVUAS 4586.2 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.293 | 1.958 | 1 | 0 | 0
| 50 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 4586_2_E12.0 |  | CVUAS 4586.2 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.339 | 1.838 | 1 | 0 | 0
# | 51 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 86_2_B4.0 |  | CVUAS 86.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.991 | 2.19 | 1 | 0 | 0
| 51 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 86_2_B5.0 |  | CVUAS 86.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.972 | 2.539 | 1 | 0 | 0
| 51 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 86_2_B6.0 |  | CVUAS 86.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.991 | 2.501 | 1 | 0 | 0
| 51 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 86_2_B4.0 |  | CVUAS 86.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.987 | 1.92 | 1 | 0 | 0
| 51 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 86_2_B5.0 |  | CVUAS 86.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.986 | 2.113 | 1 | 0 | 0
| 51 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 86_2_B6.0 |  | CVUAS 86.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.987 | 2.194 | 1 | 0 | 0
# | 52 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5453_2_H7.0 |  | CVUAS 5453.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.992 | 2.162 | 1 | 0 | 0
| 52 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5453_2_H8.0 |  | CVUAS 5453.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.995 | 2.224 | 1 | 0 | 0
| 52 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5453_2_H9.0 |  | CVUAS 5453.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.995 | 2.26 | 1 | 0 | 0
| 52 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5453_2_H7.0 |  | CVUAS 5453.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.994 | 2.342 | 1 | 0 | 0
| 52 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5453_2_H8.0 |  | CVUAS 5453.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.992 | 2.249 | 1 | 0 | 0
| 52 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5453_2_H9.0 |  | CVUAS 5453.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.993 | 2.483 | 1 | 0 | 0
# | 53 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 6844_2_C4.0 |  | CVUAS 6844.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.993 | 2.043 | 1 | 0 | 0
| 53 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 6844_2_C5.0 |  | CVUAS 6844.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.993 | 2.344 | 1 | 0 | 0
| 53 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 6844_2_C6.0 |  | CVUAS 6844.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.995 | 2.308 | 1 | 0 | 0
| 53 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 6844_2_C4.0 |  | CVUAS 6844.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.994 | 1.682 | 1 | 0 | 0
| 53 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 6844_2_C5.0 |  | CVUAS 6844.2 | FALSE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.976 | 2.752 | 1 | 0 | 0
| 53 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 6844_2_C6.0 |  | CVUAS 6844.2 | FALSE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.964 | 2.705 | 1 | 0 | 0
| 53 | 231221-1432-BayP-Sal Vali WDH_CVUAS 6844_2_A4.0 |  | CVUAS 6844.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.992 | 1.859 | 1 | 0 | 0
| 53 | 231221-1432-BayP-Sal Vali WDH_CVUAS 6844_2_A5.0 |  | CVUAS 6844.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.988 | 2.257 | 1 | 0 | 0
| 53 | 231221-1432-BayP-Sal Vali WDH_CVUAS 6844_2_A6.0 |  | CVUAS 6844.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.99 | 2.128 | 1 | 0 | 0
# | 54 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4446_2_A4.0 |  | CVUAS 4446.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.989 | 2.039 | 1 | 0 | 0
| 54 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4446_2_A5.0 |  | CVUAS 4446.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.989 | 1.867 | 1 | 0 | 0
| 54 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4446_2_A6.0 |  | CVUAS 4446.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.991 | 2.136 | 1 | 0 | 0
| 54 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4446_2_A4.0 |  | CVUAS 4446.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.99 | 2.055 | 1 | 0 | 0
| 54 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4446_2_A5.0 |  | CVUAS 4446.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.994 | 2.223 | 1 | 0 | 0
| 54 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4446_2_A6.0 |  | CVUAS 4446.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.992 | 2.199 | 1 | 0 | 0
# | 55 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 1070 _D7.0 |  | CVUAS 1070 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.939 | 2.82 | 1 | 0 | 0
| 55 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 1070 _D8.0 |  | CVUAS 1070 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.964 | 2.975 | 1 | 0 | 0
| 55 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 1070 _D9.0 |  | CVUAS 1070 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.926 | 3.195 | 1 | 0 | 0
| 55 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 1070 _D7.0 |  | CVUAS 1070 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.953 | 3.066 | 1 | 0 | 0
| 55 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 1070 _D8.0 |  | CVUAS 1070 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.948 | 3.301 | 1 | 0 | 0
| 55 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 1070 _D9.0 |  | CVUAS 1070 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.941 | 3.388 | 1 | 0 | 0
# | 56 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11190_A7.0 |  | CVUAS 11190 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.93 | 3.235 | 1 | 0 | 0
| 56 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11190_A8.0 |  | CVUAS 11190 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.807 | 3.232 | 1 | 0 | 0
| 56 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11190_A9.0 |  | CVUAS 11190 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.943 | 2.834 | 1 | 0 | 0
| 56 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11190_A7.0 |  | CVUAS 11190 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.949 | 3.183 | 1 | 0 | 0
| 56 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11190_A8.0 |  | CVUAS 11190 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.874 | 3.331 | 1 | 0 | 0
| 56 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11190_A9.0 |  | CVUAS 11190 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.962 | 2.976 | 1 | 0 | 0
# | 57 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11201_C7.0 |  | CVUAS 11201 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.932 | 3.029 | 1 | 0 | 0
| 57 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11201_C8.0 |  | CVUAS 11201 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.956 | 3.097 | 1 | 0 | 0
| 57 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11201_C9.0 |  | CVUAS 11201 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.37 | 3.376 | 1 | 0 | 0
| 57 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11201_C7.0 |  | CVUAS 11201 | FALSE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.645 | 3.625 | 0 | 1 | 0
| 57 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11201_C8.0 |  | CVUAS 11201 | FALSE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.591 | 3.729 | 0 | 1 | 0
| 57 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11201_C9.0 |  | CVUAS 11201 | FALSE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.637 | 3.769 | 0 | 1 | 0
| 57 | 231221-1432-BayP-Sal Vali WDH_CVUAS 11201_F4.0 |  | CVUAS 11201 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.877 | 2.783 | 1 | 0 | 0
| 57 | 231221-1432-BayP-Sal Vali WDH_CVUAS 11201_F5.0 |  | CVUAS 11201 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.825 | 3.584 | 0 | 1 | 0
| 57 | 231221-1432-BayP-Sal Vali WDH_CVUAS 11201_F6.0 |  | CVUAS 11201 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.922 | 3.152 | 1 | 0 | 0
# | 58 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31155_D4.0 |  | CVUAS 31155 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.704 | 2.606 | 1 | 0 | 0
| 58 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31155_D5.0 |  | CVUAS 31155 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.851 | 2.591 | 1 | 0 | 0
| 58 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31155_D6.0 |  | CVUAS 31155 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.928 | 2.61 | 1 | 0 | 0
| 58 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31155_D4.0 |  | CVUAS 31155 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.939 | 2.67 | 1 | 0 | 0
| 58 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31155_D5.0 |  | CVUAS 31155 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.957 | 2.658 | 1 | 0 | 0
| 58 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31155_D6.0 |  | CVUAS 31155 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.944 | 2.859 | 1 | 0 | 0
# | 59 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 31255_D10.0 |  | CVUAS 31255 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.861 | 3.101 | 1 | 0 | 0
| 59 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 31255_D11.0 |  | CVUAS 31255 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.825 | 3.017 | 1 | 0 | 0
| 59 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 31255_D12.0 |  | CVUAS 31255 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.813 | 3.158 | 1 | 0 | 0
| 59 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 31255_H10.0 |  | CVUAS 31255 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.751 | 2.837 | 1 | 0 | 0
| 59 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 31255_H11.0 |  | CVUAS 31255 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.914 | 2.928 | 1 | 0 | 0
| 59 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 31255_H12.0 |  | CVUAS 31255 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.865 | 2.966 | 1 | 0 | 0
# | 60 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9531_F10.0 |  | CVUAS 9531 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.791 | 3.072 | 1 | 0 | 0
| 60 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9531_F11.0 |  | CVUAS 9531 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.754 | 3.12 | 1 | 0 | 0
| 60 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9531_F12.0 |  | CVUAS 9531 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.761 | 3.092 | 1 | 0 | 0
| 60 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9531_F10.0 |  | CVUAS 9531 | FALSE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.949 | 3.059 | 1 | 0 | 0
| 60 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9531_F11.0 |  | CVUAS 9531 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.818 | 3.252 | 1 | 0 | 0
| 60 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9531_F12.0 |  | CVUAS 9531 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.923 | 3.309 | 1 | 0 | 0
# | 61 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 1148_2_F7.0 |  | CVUAS 1148.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.995 | 2.124 | 1 | 0 | 0
| 61 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 1148_2_F8.0 |  | CVUAS 1148.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.997 | 2.239 | 1 | 0 | 0
| 61 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 1148_2_F9.0 |  | CVUAS 1148.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.994 | 2.428 | 1 | 0 | 0
| 61 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 1148_2_F7.0 |  | CVUAS 1148.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.996 | 2.551 | 1 | 0 | 0
| 61 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 1148_2_F8.0 |  | CVUAS 1148.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.993 | 2.757 | 1 | 0 | 0
| 61 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 1148_2_F9.0 |  | CVUAS 1148.2 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.993 | 2.633 | 1 | 0 | 0
# | 62 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33265_D4.0 |  | CVUAS 33265 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.99 | 2.395 | 1 | 0 | 0
| 62 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33265_D5.0 |  | CVUAS 33265 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.991 | 2.619 | 1 | 0 | 0
| 62 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33265_D6.0 |  | CVUAS 33265 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.988 | 2.623 | 1 | 0 | 0
| 62 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33265_D4.0 |  | CVUAS 33265 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.992 | 2.098 | 1 | 0 | 0
| 62 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33265_D5.0 |  | CVUAS 33265 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.99 | 1.994 | 1 | 0 | 0
| 62 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33265_D6.0 |  | CVUAS 33265 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.988 | 2.103 | 1 | 0 | 0
# | 63 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31154_C4.0 |  | CVUAS 31154 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.988 | 2.338 | 1 | 0 | 0
| 63 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31154_C5.0 |  | CVUAS 31154 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.983 | 2.206 | 1 | 0 | 0
| 63 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31154_C6.0 |  | CVUAS 31154 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.985 | 2.157 | 1 | 0 | 0
| 63 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31154_C4.0 |  | CVUAS 31154 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.966 | 2.042 | 1 | 0 | 0
| 63 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31154_C5.0 |  | CVUAS 31154 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.945 | 2.667 | 1 | 0 | 0
| 63 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31154_C6.0 |  | CVUAS 31154 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.96 | 2.391 | 1 | 0 | 0
# | 64 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32815_E10.0 |  | CVUAS 32815 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.997 | 1.85 | 1 | 0 | 0
| 64 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32815_E11.0 |  | CVUAS 32815 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.995 | 2.053 | 1 | 0 | 0
| 64 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32815_E12.0 |  | CVUAS 32815 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.993 | 2.372 | 1 | 0 | 0
| 64 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32815_E10.0 |  | CVUAS 32815 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.994 | 2.271 | 1 | 0 | 0
| 64 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32815_E11.0 |  | CVUAS 32815 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.994 | 2.295 | 1 | 0 | 0
| 64 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32815_E12.0 |  | CVUAS 32815 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.993 | 2.312 | 1 | 0 | 0
# | 65 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32816_F10.0 |  | CVUAS 32816 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.989 | 2.514 | 1 | 0 | 0
| 65 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32816_F11.0 |  | CVUAS 32816 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.991 | 2.489 | 1 | 0 | 0
| 65 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32816_F12.0 |  | CVUAS 32816 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.99 | 2.557 | 1 | 0 | 0
| 65 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32816_F10.0 |  | CVUAS 32816 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.977 | 2.753 | 1 | 0 | 0
| 65 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32816_F11.0 |  | CVUAS 32816 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.974 | 2.666 | 1 | 0 | 0
| 65 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32816_F12.0 |  | CVUAS 32816 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.978 | 2.618 | 1 | 0 | 0
# | 66 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34932_C10.0 |  | CVUAS 34932 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.964 | 2.773 | 1 | 0 | 0
| 66 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34932_C11.0 |  | CVUAS 34932 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.977 | 2.687 | 1 | 0 | 0
| 66 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34932_C12.0 |  | CVUAS 34932 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.936 | 2.871 | 1 | 0 | 0
| 66 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34932_C10.0 |  | CVUAS 34932 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.983 | 2.735 | 1 | 0 | 0
| 66 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34932_C11.0 |  | CVUAS 34932 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.96 | 3.008 | 1 | 0 | 0
| 66 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34932_C12.0 |  | CVUAS 34932 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.975 | 2.857 | 1 | 0 | 0
# | 67 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34971_F10.0 |  | CVUAS 34971 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.685 | 3.401 | 1 | 0 | 0
| 67 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34971_F11.0 |  | CVUAS 34971 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.879 | 2.929 | 1 | 0 | 0
| 67 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34971_F12.0 |  | CVUAS 34971 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.874 | 2.755 | 1 | 0 | 0
| 67 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34971_F10.0 |  | CVUAS 34971 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.978 | 2.412 | 1 | 0 | 0
| 67 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34971_F11.0 |  | CVUAS 34971 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.878 | 2.803 | 1 | 0 | 0
| 67 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34971_F12.0 |  | CVUAS 34971 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.943 | 2.365 | 1 | 0 | 0
# | 68 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33992_F7.0 |  | CVUAS 33992 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.994 | 2.876 | 1 | 0 | 0
| 68 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33992_F8.0 |  | CVUAS 33992 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.993 | 2.779 | 1 | 0 | 0
| 68 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33992_F9.0 |  | CVUAS 33992 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.995 | 2.352 | 1 | 0 | 0
| 68 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33992_F7.0 |  | CVUAS 33992 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.994 | 2.278 | 1 | 0 | 0
| 68 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33992_F8.0 |  | CVUAS 33992 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.997 | 2.357 | 1 | 0 | 0
| 68 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33992_F9.0 |  | CVUAS 33992 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.993 | 2.493 | 1 | 0 | 0
# | 69 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31528_A7.0 |  | CVUAS 31528 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.992 | 2.409 | 1 | 0 | 0
| 69 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31528_A8.0 |  | CVUAS 31528 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.991 | 2.315 | 1 | 0 | 0
| 69 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 31528_A9.0 |  | CVUAS 31528 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.99 | 2.213 | 1 | 0 | 0
| 69 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31528_A7.0 |  | CVUAS 31528 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.994 | 2.247 | 1 | 0 | 0
| 69 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31528_A8.0 |  | CVUAS 31528 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.994 | 2.062 | 1 | 0 | 0
| 69 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 31528_A9.0 |  | CVUAS 31528 | TRUE | non - O:7 (C1-C4) | O:8 (C2-C3) | 1 | 0.994 | 2.167 | 1 | 0 | 0
# | 70 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11556_2_B10.0 |  | CVUAS 11556.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.251 | 1 | 0 | 0
| 70 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11556_2_B11.0 |  | CVUAS 11556.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.156 | 1 | 0 | 0
| 70 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11556_2_B12.0 |  | CVUAS 11556.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.106 | 1 | 0 | 0
| 70 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11556_2_B10.0 |  | CVUAS 11556.2 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.902 | 2.046 | 1 | 0 | 0
| 70 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11556_2_B11.0 |  | CVUAS 11556.2 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.98 | 2.349 | 1 | 0 | 0
| 70 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11556_2_B12.0 |  | CVUAS 11556.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.995 | 2.375 | 1 | 0 | 0
| 70 | 231221-1432-BayP-Sal Vali WDH_CVUAS 11556_2_A7.0 |  | CVUAS 11556.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.999 | 2.221 | 1 | 0 | 0
| 70 | 231221-1432-BayP-Sal Vali WDH_CVUAS 11556_2_A8.0 |  | CVUAS 11556.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.999 | 2.265 | 1 | 0 | 0
| 70 | 231221-1432-BayP-Sal Vali WDH_CVUAS 11556_2_A9.0 |  | CVUAS 11556.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.328 | 1 | 0 | 0
# | 71 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7670_2_C7.0 |  | CVUAS 7670.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.165 | 1 | 0 | 0
| 71 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7670_2_C8.0 |  | CVUAS 7670.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.102 | 1 | 0 | 0
| 71 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7670_2_C9.0 |  | CVUAS 7670.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.994 | 2.05 | 1 | 0 | 0
| 71 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7670_2_C7.0 |  | CVUAS 7670.2 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.887 | 2.782 | 1 | 0 | 0
| 71 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7670_2_C8.0 |  | CVUAS 7670.2 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.858 | 2.772 | 1 | 0 | 0
| 71 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7670_2_C9.0 |  | CVUAS 7670.2 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.862 | 2.949 | 1 | 0 | 0
| 71 | 231221-1432-BayP-Sal Vali WDH_CVUAS 7670_2_C4.0 |  | CVUAS 7670.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.991 | 1.664 | 1 | 0 | 0
| 71 | 231221-1432-BayP-Sal Vali WDH_CVUAS 7670_2_C5.0 |  | CVUAS 7670.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.992 | 1.787 | 1 | 0 | 0
| 71 | 231221-1432-BayP-Sal Vali WDH_CVUAS 7670_2_C6.0 |  | CVUAS 7670.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.979 | 1.669 | 1 | 0 | 0
# | 72 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9544_2_G10.0 |  | CVUAS 9544.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.399 | 1 | 0 | 0
| 72 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9544_2_G11.0 |  | CVUAS 9544.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 2.301 | 1 | 0 | 0
| 72 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9544_2_G12.0 |  | CVUAS 9544.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 2.36 | 1 | 0 | 0
| 72 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9544_2_G10.0 |  | CVUAS 9544.2 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.986 | 2.102 | 1 | 0 | 0
| 72 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9544_2_G11.0 |  | CVUAS 9544.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.991 | 2.217 | 1 | 0 | 0
| 72 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9544_2_G12.0 |  | CVUAS 9544.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.99 | 2.04 | 1 | 0 | 0
# | 73 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11434_F7.0 |  | CVUAS 11434 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.347 | 1 | 0 | 0
| 73 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11434_F8.0 |  | CVUAS 11434 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 1.857 | 1 | 0 | 0
| 73 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11434_F9.0 |  | CVUAS 11434 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.03 | 1 | 0 | 0
| 73 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11434_F7.0 |  | CVUAS 11434 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.994 | 2.121 | 1 | 0 | 0
| 73 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11434_F8.0 |  | CVUAS 11434 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.993 | 2.031 | 1 | 0 | 0
| 73 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11434_F9.0 |  | CVUAS 11434 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.982 | 1.356 | 1 | 0 | 0
# | 74 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 30458_E10.0 |  | CVUAS 30458 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.985 | 1.483 | 1 | 0 | 0
| 74 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 30458_E11.0 |  | CVUAS 30458 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.989 | 1.631 | 1 | 0 | 0
| 74 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 30458_E12.0 |  | CVUAS 30458 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.294 | 1 | 0 | 0
| 74 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 30458_E10.0 |  | CVUAS 30458 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.995 | 2.235 | 1 | 0 | 0
| 74 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 30458_E11.0 |  | CVUAS 30458 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.977 | 1.621 | 1 | 0 | 0
| 74 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 30458_E12.0 |  | CVUAS 30458 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 2.256 | 1 | 0 | 0
# | 75 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11141_2_H4.0 |  | CVUAS 11141.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.992 | 2.211 | 1 | 0 | 0
| 75 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11141_2_H5.0 |  | CVUAS 11141.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.993 | 2.139 | 1 | 0 | 0
| 75 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11141_2_H6.0 |  | CVUAS 11141.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.994 | 1.77 | 1 | 0 | 0
| 75 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11141_2_H4.0 |  | CVUAS 11141.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.986 | 2.224 | 1 | 0 | 0
| 75 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11141_2_H5.0 |  | CVUAS 11141.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.985 | 2.151 | 1 | 0 | 0
| 75 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11141_2_H6.0 |  | CVUAS 11141.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.989 | 2.129 | 1 | 0 | 0
# | 76 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33131_B4.0 |  | CVUAS 33131 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 1.988 | 1 | 0 | 0
| 76 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33131_B5.0 |  | CVUAS 33131 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.269 | 1 | 0 | 0
| 76 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33131_B6.0 |  | CVUAS 33131 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.995 | 2.152 | 1 | 0 | 0
| 76 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33131_B4.0 |  | CVUAS 33131 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.978 | 1.813 | 1 | 0 | 0
| 76 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33131_B5.0 |  | CVUAS 33131 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.789 | 2.653 | 1 | 0 | 0
| 76 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33131_B6.0 |  | CVUAS 33131 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.962 | 1.786 | 1 | 0 | 0
# | 77 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33274_E4.0 |  | CVUAS 33274 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.995 | 1.855 | 1 | 0 | 0
| 77 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33274_E5.0 |  | CVUAS 33274 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.004 | 1 | 0 | 0
| 77 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33274_E6.0 |  | CVUAS 33274 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 1.988 | 1 | 0 | 0
| 77 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33274_E4.0 |  | CVUAS 33274 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.971 | 2.11 | 1 | 0 | 0
| 77 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33274_E5.0 |  | CVUAS 33274 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.949 | 1.858 | 1 | 0 | 0
| 77 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33274_E6.0 |  | CVUAS 33274 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.986 | 1.889 | 1 | 0 | 0
# | 78 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34824_H7.0 |  | CVUAS 34824 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.999 | 2.279 | 1 | 0 | 0
| 78 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34824_H8.0 |  | CVUAS 34824 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.999 | 2.102 | 1 | 0 | 0
| 78 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34824_H9.0 |  | CVUAS 34824 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.053 | 1 | 0 | 0
| 78 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34824_H7.0 |  | CVUAS 34824 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.145 | 1 | 0 | 0
| 78 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34824_H8.0 |  | CVUAS 34824 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.394 | 1 | 0 | 0
| 78 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34824_H9.0 |  | CVUAS 34824 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 2.133 | 1 | 0 | 0
# | 79 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34826_A10.0 |  | CVUAS 34826 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.681 | 1 | 0 | 0
| 79 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34826_A11.0 |  | CVUAS 34826 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.853 | 1 | 0 | 0
| 79 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 34826_A12.0 |  | CVUAS 34826 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.832 | 1 | 0 | 0
| 79 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34826_A10.0 |  | CVUAS 34826 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.999 | 2.699 | 1 | 0 | 0
| 79 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34826_A11.0 |  | CVUAS 34826 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.655 | 1 | 0 | 0
| 79 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 34826_A12.0 |  | CVUAS 34826 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.999 | 2.708 | 1 | 0 | 0
# | 80 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35000_C4.0 |  | CVUAS 35000 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.116 | 1 | 0 | 0
| 80 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35000_C5.0 |  | CVUAS 35000 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.09 | 1 | 0 | 0
| 80 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35000_C6.0 |  | CVUAS 35000 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.2 | 1 | 0 | 0
| 80 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35000_C4.0 |  | CVUAS 35000 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.99 | 2.022 | 1 | 0 | 0
| 80 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35000_C5.0 |  | CVUAS 35000 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.99 | 2.034 | 1 | 0 | 0
| 80 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35000_C6.0 |  | CVUAS 35000 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.991 | 2.119 | 1 | 0 | 0
# | 81 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35070_B7.0 |  | CVUAS 35070 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.605 | 1 | 0 | 0
| 81 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35070_B8.0 |  | CVUAS 35070 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.584 | 1 | 0 | 0
| 81 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35070_B9.0 |  | CVUAS 35070 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.53 | 1 | 0 | 0
| 81 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35070_B7.0 |  | CVUAS 35070 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.992 | 1.944 | 1 | 0 | 0
| 81 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35070_B8.0 |  | CVUAS 35070 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.992 | 1.975 | 1 | 0 | 0
| 81 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35070_B9.0 |  | CVUAS 35070 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.991 | 1.923 | 1 | 0 | 0
# | 82 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35122_F7.0 |  | CVUAS 35122 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.982 | 1.829 | 1 | 0 | 0
| 82 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35122_F8.0 |  | CVUAS 35122 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.984 | 1.76 | 1 | 0 | 0
| 82 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35122_F9.0 |  | CVUAS 35122 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.981 | 2.014 | 1 | 0 | 0
| 82 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35122_F7.0 |  | CVUAS 35122 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.705 | 1.478 | 1 | 0 | 0
| 82 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35122_F8.0 |  | CVUAS 35122 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.952 | 1.783 | 1 | 0 | 0
| 82 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35122_F9.0 |  | CVUAS 35122 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.943 | 1.77 | 1 | 0 | 0
# | 83 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35135_G7.0 |  | CVUAS 35135 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.444 | 1 | 0 | 0
| 83 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35135_G8.0 |  | CVUAS 35135 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.374 | 1 | 0 | 0
| 83 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35135_G9.0 |  | CVUAS 35135 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.382 | 1 | 0 | 0
| 83 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35135_G7.0 |  | CVUAS 35135 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.387 | 1 | 0 | 0
| 83 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35135_G8.0 |  | CVUAS 35135 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.391 | 1 | 0 | 0
| 83 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35135_G9.0 |  | CVUAS 35135 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.995 | 2.324 | 1 | 0 | 0
# | 84 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35176_H7.0 |  | CVUAS 35176 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 1.942 | 1 | 0 | 0
| 84 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35176_H8.0 |  | CVUAS 35176 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 1.937 | 1 | 0 | 0
| 84 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35176_H9.0 |  | CVUAS 35176 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 2.096 | 1 | 0 | 0
| 84 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35176_H7.0 |  | CVUAS 35176 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.366 | 1 | 0 | 0
| 84 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35176_H8.0 |  | CVUAS 35176 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.995 | 2.225 | 1 | 0 | 0
| 84 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35176_H9.0 |  | CVUAS 35176 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.994 | 2.317 | 1 | 0 | 0
# | 85 | 240112-1231-nov_dyk-sal vali_CVUAS 11284_G10.0 |  | CVUAS 11284 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.096 | 1 | 0 | 0
| 85 | 240112-1231-nov_dyk-sal vali_CVUAS 11284_G11.0 |  | CVUAS 11284 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.129 | 1 | 0 | 0
| 85 | 240112-1231-nov_dyk-sal vali_CVUAS 11284_G12.0 |  | CVUAS 11284 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.042 | 1 | 0 | 0
| 85 | 240112-1231-nov_dyk-sal vali_CVUAS 11284_G7.0 |  | CVUAS 11284 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 2.152 | 1 | 0 | 0
| 85 | 240112-1231-nov_dyk-sal vali_CVUAS 11284_G8.0 |  | CVUAS 11284 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.139 | 1 | 0 | 0
| 85 | 240112-1231-nov_dyk-sal vali_CVUAS 11284_G9.0 |  | CVUAS 11284 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 2.13 | 1 | 0 | 0
# | 86 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11298_3_D7.0 |  | CVUAS 11298.3 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.992 | 1.995 | 1 | 0 | 0
| 86 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11298_3_D8.0 |  | CVUAS 11298.3 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 2.091 | 1 | 0 | 0
| 86 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11298_3_D9.0 |  | CVUAS 11298.3 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.993 | 2.082 | 1 | 0 | 0
| 86 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11298_3_D7.0 |  | CVUAS 11298.3 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.889 | 2.546 | 1 | 0 | 0
| 86 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11298_3_D8.0 |  | CVUAS 11298.3 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.972 | 2.654 | 1 | 0 | 0
| 86 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11298_3_D9.0 |  | CVUAS 11298.3 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.979 | 2.583 | 1 | 0 | 0
| 86 | 231221-1432-BayP-Sal Vali WDH_CVUAS 11298_3_G4.0 |  | CVUAS 11298.3 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.994 | 2.244 | 1 | 0 | 0
| 86 | 231221-1432-BayP-Sal Vali WDH_CVUAS 11298_3_G5.0 |  | CVUAS 11298.3 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.959 | 2.64 | 1 | 0 | 0
| 86 | 231221-1432-BayP-Sal Vali WDH_CVUAS 11298_3_G6.0 |  | CVUAS 11298.3 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.931 | 2.519 | 1 | 0 | 0
# | 87 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32489_H7.0 |  | CVUAS 32489 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.688 | 1 | 0 | 0
| 87 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32489_H8.0 |  | CVUAS 32489 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.999 | 2.694 | 1 | 0 | 0
| 87 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32489_H9.0 |  | CVUAS 32489 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 3.117 | 1 | 0 | 0
| 87 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32489_H7.0 |  | CVUAS 32489 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.84 | 1 | 0 | 0
| 87 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32489_H8.0 |  | CVUAS 32489 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 2.92 | 1 | 0 | 0
| 87 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32489_H9.0 |  | CVUAS 32489 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.851 | 1 | 0 | 0
# | 88 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7511_A7.0 |  | CVUAS 7511 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.548 | 1 | 0 | 0
| 88 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7511_A8.0 |  | CVUAS 7511 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.329 | 1 | 0 | 0
| 88 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7511_A9.0 |  | CVUAS 7511 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.994 | 2.404 | 1 | 0 | 0
| 88 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7511_A7.0 |  | CVUAS 7511 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.589 | 1 | 0 | 0
| 88 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7511_A8.0 |  | CVUAS 7511 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 2.327 | 1 | 0 | 0
| 88 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7511_A9.0 |  | CVUAS 7511 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.994 | 2.08 | 1 | 0 | 0
# | 89 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5362_2_G7.0 |  | CVUAS 5362.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.362 | 1 | 0 | 0
| 89 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5362_2_G8.0 |  | CVUAS 5362.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 2.556 | 1 | 0 | 0
| 89 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5362_2_G9.0 |  | CVUAS 5362.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.483 | 1 | 0 | 0
| 89 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5362_2_G7.0 |  | CVUAS 5362.2 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 1.533 | 1 | 0 | 0
| 89 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5362_2_G8.0 |  | CVUAS 5362.2 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.995 | 1.652 | 1 | 0 | 0
| 89 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5362_2_G9.0 |  | CVUAS 5362.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 1.821 | 1 | 0 | 0
# | 90 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 8782_B10.0 |  | CVUAS 8782 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.946 | 2.83 | 1 | 0 | 0
| 90 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 8782_B11.0 |  | CVUAS 8782 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.912 | 2.735 | 1 | 0 | 0
| 90 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 8782_B12.0 |  | CVUAS 8782 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.888 | 2.765 | 1 | 0 | 0
| 90 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 8782_B10.0 |  | CVUAS 8782 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.637 | 3.017 | 1 | 0 | 0
| 90 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 8782_B11.0 |  | CVUAS 8782 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.387 | 2.641 | 1 | 0 | 0
| 90 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 8782_B12.0 |  | CVUAS 8782 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.587 | 2.758 | 1 | 0 | 0
# | 91 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 3238_2_E10.0 |  | CVUAS 3238.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 2.025 | 1 | 0 | 0
| 91 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 3238_2_E11.0 |  | CVUAS 3238.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 2.001 | 1 | 0 | 0
| 91 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 3238_2_E12.0 |  | CVUAS 3238.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 2.144 | 1 | 0 | 0
| 91 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 3238_2_E10.0 |  | CVUAS 3238.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.999 | 2.092 | 1 | 0 | 0
| 91 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 3238_2_E11.0 |  | CVUAS 3238.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.232 | 1 | 0 | 0
| 91 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 3238_2_E12.0 |  | CVUAS 3238.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.998 | 2.261 | 1 | 0 | 0
# | 92 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 2553_2 _C10.0 |  | CVUAS 2553.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.209 | 1 | 0 | 0
| 92 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 2553_2 _C11.0 |  | CVUAS 2553.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.997 | 2.337 | 1 | 0 | 0
| 92 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 2553_2 _C12.0 |  | CVUAS 2553.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.34 | 1 | 0 | 0
| 92 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 2553_2 _C10.0 |  | CVUAS 2553.2 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.366 | 1 | 0 | 0
| 92 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 2553_2 _C11.0 |  | CVUAS 2553.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.996 | 2.443 | 1 | 0 | 0
| 92 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 2553_2 _C12.0 |  | CVUAS 2553.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.995 | 2.449 | 1 | 0 | 0
# | 93 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 122_2_C4.0 |  | CVUAS 122.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.906 | 1.833 | 1 | 0 | 0
| 93 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 122_2_C5.0 |  | CVUAS 122.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.899 | 1.842 | 1 | 0 | 0
| 93 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 122_2_C6.0 |  | CVUAS 122.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.897 | 1.838 | 1 | 0 | 0
| 93 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 122_2_C4.0 |  | CVUAS 122.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.949 | 1.63 | 1 | 0 | 0
| 93 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 122_2_C5.0 |  | CVUAS 122.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.951 | 1.847 | 1 | 0 | 0
| 93 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 122_2_C6.0 |  | CVUAS 122.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.957 | 2.015 | 1 | 0 | 0
# | 94 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 30311_2_D10.0 |  | CVUAS 30311.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.994 | 1.572 | 1 | 0 | 0
| 94 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 30311_2_D11.0 |  | CVUAS 30311.2 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.984 | 1.407 | 1 | 0 | 0
| 94 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 30311_2_D12.0 |  | CVUAS 30311.2 | FALSE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.986 | 1.431 | 1 | 0 | 0
| 94 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 30311_2_D10.0 |  | CVUAS 30311.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.974 | 1.575 | 1 | 0 | 0
| 94 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 30311_2_D11.0 |  | CVUAS 30311.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.986 | 2.48 | 1 | 0 | 0
| 94 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 30311_2_D12.0 |  | CVUAS 30311.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.99 | 2.728 | 1 | 0 | 0
# | 95 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5560_C10.0 |  | CVUAS 5560 | TRUE | non - O:7 (C1-C4) | O:9,46 (D2) | 1 | 0.971 | 2.38 | 1 | 0 | 0
| 95 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5560_C11.0 |  | CVUAS 5560 | TRUE | non - O:7 (C1-C4) | O:9,46 (D2) | 1 | 0.969 | 2.449 | 1 | 0 | 0
| 95 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5560_C12.0 |  | CVUAS 5560 | TRUE | non - O:7 (C1-C4) | O:9,46 (D2) | 1 | 0.972 | 2.462 | 1 | 0 | 0
| 95 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5560_C10.0 |  | CVUAS 5560 | TRUE | non - O:7 (C1-C4) | O:9,46 (D2) | 1 | 0.967 | 2.153 | 1 | 0 | 0
| 95 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5560_C11.0 |  | CVUAS 5560 | TRUE | non - O:7 (C1-C4) | O:9,46 (D2) | 1 | 0.965 | 2.218 | 1 | 0 | 0
| 95 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5560_C12.0 |  | CVUAS 5560 | TRUE | non - O:7 (C1-C4) | O:9,46 (D2) | 1 | 0.965 | 2.169 | 1 | 0 | 0
# | 96 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4815_G4.0 |  | CVUAS 4815 | TRUE | non - O:7 (C1-C4) | O:9,46 (D2) | 1 | 0.837 | 2.908 | 1 | 0 | 0
| 96 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4815_G5.0 |  | CVUAS 4815 | TRUE | non - O:7 (C1-C4) | O:9,46 (D2) | 1 | 0.839 | 2.954 | 1 | 0 | 0
| 96 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4815_G6.0 |  | CVUAS 4815 | TRUE | non - O:7 (C1-C4) | O:9,46 (D2) | 1 | 0.844 | 3.047 | 1 | 0 | 0
| 96 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4815_G4.0 |  | CVUAS 4815 | TRUE | non - O:7 (C1-C4) | O:9,46 (D2) | 1 | 0.837 | 2.917 | 1 | 0 | 0
| 96 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4815_G5.0 |  | CVUAS 4815 | TRUE | non - O:7 (C1-C4) | O:9,46 (D2) | 1 | 0.85 | 2.745 | 1 | 0 | 0
| 96 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4815_G6.0 |  | CVUAS 4815 | TRUE | non - O:7 (C1-C4) | O:9,46 (D2) | 1 | 0.841 | 2.798 | 1 | 0 | 0
# | 97 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35012 _D4.0 |  | CVUAS 35012 | TRUE | non - O:7 (C1-C4) | O:3,10 (E1-E2-E3) | 1 | 0.987 | 2.31 | 1 | 0 | 0
| 97 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35012 _D5.0 |  | CVUAS 35012 | TRUE | non - O:7 (C1-C4) | O:3,10 (E1-E2-E3) | 1 | 0.685 | 1.931 | 1 | 0 | 0
| 97 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35012 _D6.0 |  | CVUAS 35012 | TRUE | non - O:7 (C1-C4) | O:1,3,19 (E4) | 1 | 0.527 | 2.143 | 1 | 0 | 0
| 97 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35012 _D4.0 |  | CVUAS 35012 | TRUE | non - O:7 (C1-C4) | O:3,10 (E1-E2-E3) | 1 | 0.987 | 1.712 | 1 | 0 | 0
| 97 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35012 _D5.0 |  | CVUAS 35012 | TRUE | non - O:7 (C1-C4) | O:3,10 (E1-E2-E3) | 1 | 0.983 | 1.861 | 1 | 0 | 0
| 97 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35012 _D6.0 |  | CVUAS 35012 | TRUE | non - O:7 (C1-C4) | O:3,10 (E1-E2-E3) | 1 | 0.995 | 1.905 | 1 | 0 | 0
# | 98 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35073_C7.0 |  | CVUAS 35073 | TRUE | non - O:7 (C1-C4) | O:3,10 (E1-E2-E3) | 1 | 0.994 | 2.919 | 1 | 0 | 0
| 98 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35073_C8.0 |  | CVUAS 35073 | TRUE | non - O:7 (C1-C4) | O:3,10 (E1-E2-E3) | 1 | 0.994 | 2.836 | 1 | 0 | 0
| 98 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35073_C9.0 |  | CVUAS 35073 | TRUE | non - O:7 (C1-C4) | O:3,10 (E1-E2-E3) | 1 | 0.976 | 2.873 | 1 | 0 | 0
| 98 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35073_C7.0 |  | CVUAS 35073 | TRUE | non - O:7 (C1-C4) | O:3,10 (E1-E2-E3) | 1 | 0.99 | 2.504 | 1 | 0 | 0
| 98 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35073_C8.0 |  | CVUAS 35073 | TRUE | non - O:7 (C1-C4) | O:3,10 (E1-E2-E3) | 1 | 0.993 | 2.51 | 1 | 0 | 0
| 98 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35073_C9.0 |  | CVUAS 35073 | TRUE | non - O:7 (C1-C4) | O:3,10 (E1-E2-E3) | 1 | 0.993 | 2.44 | 1 | 0 | 0
# | 99 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9446_E10.0 |  | CVUAS 9446 | TRUE | non - O:7 (C1-C4) | O:1,3,19 (E4) | 1 | 0.996 | 2.734 | 1 | 0 | 0
| 99 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9446_E11.0 |  | CVUAS 9446 | TRUE | non - O:7 (C1-C4) | O:1,3,19 (E4) | 1 | 0.996 | 2.743 | 1 | 0 | 0
| 99 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9446_E12.0 |  | CVUAS 9446 | TRUE | non - O:7 (C1-C4) | O:1,3,19 (E4) | 1 | 0.994 | 2.675 | 1 | 0 | 0
| 99 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9446_E10.0 |  | CVUAS 9446 | TRUE | non - O:7 (C1-C4) | O:1,3,19 (E4) | 1 | 0.982 | 2.135 | 1 | 0 | 0
| 99 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9446_E11.0 |  | CVUAS 9446 | TRUE | non - O:7 (C1-C4) | O:1,3,19 (E4) | 1 | 0.988 | 2.521 | 1 | 0 | 0
| 99 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9446_E12.0 |  | CVUAS 9446 | TRUE | non - O:7 (C1-C4) | O:1,3,19 (E4) | 1 | 0.992 | 2.5 | 1 | 0 | 0
# | 100 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33893_E7.0 |  | CVUAS 33893 | TRUE | non - O:7 (C1-C4) | O:1,3,19 (E4) | 1 | 0.974 | 2.386 | 1 | 0 | 0
| 100 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33893_E8.0 |  | CVUAS 33893 | TRUE | non - O:7 (C1-C4) | O:1,3,19 (E4) | 1 | 0.943 | 2.742 | 1 | 0 | 0
| 100 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33893_E9.0 |  | CVUAS 33893 | TRUE | non - O:7 (C1-C4) | O:1,3,19 (E4) | 1 | 0.989 | 2.465 | 1 | 0 | 0
| 100 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33893_E7.0 |  | CVUAS 33893 | TRUE | non - O:7 (C1-C4) | O:1,3,19 (E4) | 1 | 0.978 | 2.005 | 1 | 0 | 0
| 100 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33893_E8.0 |  | CVUAS 33893 | TRUE | non - O:7 (C1-C4) | O:1,3,19 (E4) | 1 | 0.953 | 2.115 | 1 | 0 | 0
| 100 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33893_E9.0 |  | CVUAS 33893 | TRUE | non - O:7 (C1-C4) | O:1,3,19 (E4) | 1 | 0.97 | 2.136 | 1 | 0 | 0
# | 101 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 30550_G10.0 |  | CVUAS 30550 | TRUE | non - O:7 (C1-C4) | O:11 (F) | 1 | 0.813 | 3.242 | 1 | 0 | 0
| 101 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 30550_G11.0 |  | CVUAS 30550 | TRUE | non - O:7 (C1-C4) | O:11 (F) | 1 | 0.768 | 3.28 | 1 | 0 | 0
| 101 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 30550_G12.0 |  | CVUAS 30550 | TRUE | non - O:7 (C1-C4) | O:11 (F) | 1 | 0.602 | 3.211 | 1 | 0 | 0
| 101 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 30550_G10.0 |  | CVUAS 30550 | TRUE | non - O:7 (C1-C4) | O:11 (F) | 1 | 0.352 | 3.081 | 1 | 0 | 0
| 101 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 30550_G11.0 |  | CVUAS 30550 | TRUE | non - O:7 (C1-C4) | O:11 (F) | 1 | 0.313 | 3.163 | 1 | 0 | 0
| 101 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 30550_G12.0 |  | CVUAS 30550 | TRUE | non - O:7 (C1-C4) | O:11 (F) | 1 | 0.369 | 3.26 | 1 | 0 | 0
# | 102 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 369_4_E4.0 |  | CVUAS 369.4 | TRUE | non - O:7 (C1-C4) | O:11 (F) | 1 | 0.37 | 2.439 | 1 | 0 | 0
| 102 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 369_4_E5.0 |  | CVUAS 369.4 | TRUE | non - O:7 (C1-C4) | O:11 (F) | 1 | 0.315 | 2.459 | 1 | 0 | 0
| 102 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 369_4_E6.0 |  | CVUAS 369.4 | TRUE | non - O:7 (C1-C4) | O:11 (F) | 1 | 0.304 | 2.566 | 1 | 0 | 0
| 102 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 369_4_E4.0 |  | CVUAS 369.4 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.281 | 2.175 | 1 | 0 | 0
| 102 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 369_4_E5.0 |  | CVUAS 369.4 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.574 | 2.124 | 1 | 0 | 0
| 102 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 369_4_E6.0 |  | CVUAS 369.4 | TRUE | non - O:7 (C1-C4) | O:11 (F) | 1 | 0.2 | 2.333 | 1 | 0 | 0
# | 103 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33780_B7.0 |  | CVUAS 33780 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.989 | 2.947 | 1 | 0 | 0
| 103 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33780_B8.0 |  | CVUAS 33780 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.993 | 2.834 | 1 | 0 | 0
| 103 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33780_B9.0 |  | CVUAS 33780 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.989 | 2.868 | 1 | 0 | 0
| 103 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33780_B7.0 |  | CVUAS 33780 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.96 | 2.447 | 1 | 0 | 0
| 103 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33780_B8.0 |  | CVUAS 33780 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.974 | 2.373 | 1 | 0 | 0
| 103 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33780_B9.0 |  | CVUAS 33780 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.982 | 2.552 | 1 | 0 | 0
# | 104 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5270_3_F7.0 |  | CVUAS 5270.3 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.994 | 2.649 | 1 | 0 | 0
| 104 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5270_3_F8.0 |  | CVUAS 5270.3 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.992 | 2.327 | 1 | 0 | 0
| 104 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5270_3_F9.0 |  | CVUAS 5270.3 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.981 | 2.103 | 1 | 0 | 0
| 104 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5270_3_F7.0 |  | CVUAS 5270.3 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.976 | 2.464 | 1 | 0 | 0
| 104 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5270_3_F8.0 |  | CVUAS 5270.3 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.968 | 2.515 | 1 | 0 | 0
| 104 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5270_3_F9.0 |  | CVUAS 5270.3 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.972 | 2.441 | 1 | 0 | 0
# | 105 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11555_2_A10.0 |  | CVUAS 11555.2 | TRUE | non - O:7 (C1-C4) | O:6,14 (H) | 1 | 0.604 | 2.645 | 1 | 0 | 0
| 105 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11555_2_A11.0 |  | CVUAS 11555.2 | TRUE | non - O:7 (C1-C4) | O:6,14 (H) | 1 | 0.614 | 2.646 | 1 | 0 | 0
| 105 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11555_2_A12.0 |  | CVUAS 11555.2 | TRUE | non - O:7 (C1-C4) | O:6,14 (H) | 1 | 0.631 | 2.608 | 1 | 0 | 0
| 105 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11555_2_A10.0 |  | CVUAS 11555.2 | TRUE | non - O:7 (C1-C4) | O:6,14 (H) | 1 | 0.895 | 2.608 | 1 | 0 | 0
| 105 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11555_2_A11.0 |  | CVUAS 11555.2 | TRUE | non - O:7 (C1-C4) | O:6,14 (H) | 1 | 0.959 | 2.786 | 1 | 0 | 0
| 105 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11555_2_A12.0 |  | CVUAS 11555.2 | TRUE | non - O:7 (C1-C4) | O:6,14 (H) | 1 | 0.948 | 3.051 | 1 | 0 | 0
# | 106 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4871_3_H4.0 |  | CVUAS 4871.3 | TRUE | non - O:7 (C1-C4) | O:6,14 (H) | 1 | 0.982 | 2.021 | 1 | 0 | 0
| 106 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4871_3_H5.0 |  | CVUAS 4871.3 | TRUE | non - O:7 (C1-C4) | O:6,14 (H) | 1 | 0.977 | 1.982 | 1 | 0 | 0
| 106 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4871_3_H6.0 |  | CVUAS 4871.3 | TRUE | non - O:7 (C1-C4) | O:6,14 (H) | 1 | 0.971 | 2.135 | 1 | 0 | 0
| 106 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4871_3_H4.0 |  | CVUAS 4871.3 | TRUE | non - O:7 (C1-C4) | O:6,14 (H) | 1 | 0.989 | 2.135 | 1 | 0 | 0
| 106 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4871_3_H5.0 |  | CVUAS 4871.3 | TRUE | non - O:7 (C1-C4) | O:6,14 (H) | 1 | 0.989 | 2.443 | 1 | 0 | 0
| 106 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4871_3_H6.0 |  | CVUAS 4871.3 | TRUE | non - O:7 (C1-C4) | O:6,14 (H) | 1 | 0.99 | 2.504 | 1 | 0 | 0
# | 107 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35077_E7.0 |  | CVUAS 35077 | TRUE | non - O:7 (C1-C4) | O:16 (I) | 1 | 0.982 | 3.182 | 1 | 0 | 0
| 107 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35077_E8.0 |  | CVUAS 35077 | TRUE | non - O:7 (C1-C4) | O:16 (I) | 1 | 0.973 | 3.085 | 1 | 0 | 0
| 107 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35077_E9.0 |  | CVUAS 35077 | TRUE | non - O:7 (C1-C4) | O:16 (I) | 1 | 0.979 | 3.24 | 1 | 0 | 0
| 107 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35077_E7.0 |  | CVUAS 35077 | TRUE | non - O:7 (C1-C4) | O:16 (I) | 1 | 0.624 | 2.068 | 1 | 0 | 0
| 107 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35077_E8.0 |  | CVUAS 35077 | TRUE | non - O:7 (C1-C4) | O:16 (I) | 1 | 0.862 | 2.208 | 1 | 0 | 0
| 107 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35077_E9.0 |  | CVUAS 35077 | TRUE | non - O:7 (C1-C4) | O:16 (I) | 1 | 0.849 | 2.086 | 1 | 0 | 0
# | 108 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35625_G10.0 |  | CVUAS 35625 | TRUE | non - O:7 (C1-C4) | O:16 (I) | 1 | 0.952 | 2.278 | 1 | 0 | 0
| 108 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35625_G11.0 |  | CVUAS 35625 | TRUE | non - O:7 (C1-C4) | O:16 (I) | 1 | 0.96 | 2.536 | 1 | 0 | 0
| 108 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35625_G12.0 |  | CVUAS 35625 | TRUE | non - O:7 (C1-C4) | O:16 (I) | 1 | 0.944 | 2.304 | 1 | 0 | 0
| 108 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35625_G10.0 |  | CVUAS 35625 | TRUE | non - O:7 (C1-C4) | O:16 (I) | 1 | 0.761 | 1.995 | 1 | 0 | 0
| 108 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35625_G11.0 |  | CVUAS 35625 | TRUE | non - O:7 (C1-C4) | O:16 (I) | 1 | 0.75 | 2.066 | 1 | 0 | 0
| 108 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35625_G12.0 |  | CVUAS 35625 | TRUE | non - O:7 (C1-C4) | O:16 (I) | 1 | 0.758 | 2.044 | 1 | 0 | 0
# | 109 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 3505 _G10.0 |  | CVUAS 3505 | TRUE | non - O:7 (C1-C4) | O:17 (J) | 1 | 0.329 | 2.317 | 1 | 0 | 0
| 109 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 3505 _G11.0 |  | CVUAS 3505 | TRUE | non - O:7 (C1-C4) | O:17 (J) | 1 | 0.204 | 3.097 | 1 | 0 | 0
| 109 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 3505 _G12.0 |  | CVUAS 3505 | TRUE | non - O:7 (C1-C4) | O:17 (J) | 1 | 0.295 | 2.651 | 1 | 0 | 0
| 109 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 3505 _G10.0 |  | CVUAS 3505 | TRUE | non - O:7 (C1-C4) | O:17 (J) | 1 | 0.249 | 2.712 | 1 | 0 | 0
| 109 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 3505 _G11.0 |  | CVUAS 3505 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.185 | 2.851 | 1 | 0 | 0
| 109 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 3505 _G12.0 |  | CVUAS 3505 | TRUE | non - O:7 (C1-C4) | O:17 (J) | 1 | 0.21 | 2.639 | 1 | 0 | 0
# | 110 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32901_G10.0 |  | CVUAS 32901 | TRUE | non - O:7 (C1-C4) | O:38 (P) | 1 | 0.141 | 2.401 | 1 | 0 | 0
| 110 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32901_G11.0 |  | CVUAS 32901 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.136 | 2.926 | 1 | 0 | 0
| 110 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32901_G12.0 |  | CVUAS 32901 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.115 | 2.754 | 1 | 0 | 0
| 110 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32901_G10.0 |  | CVUAS 32901 | FALSE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.133 | 3.78 | 0 | 1 | 0
| 110 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32901_G11.0 |  | CVUAS 32901 | TRUE | non - O:7 (C1-C4) | O:17 (J) | 1 | 0.202 | 2.942 | 1 | 0 | 0
| 110 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32901_G12.0 |  | CVUAS 32901 | TRUE | non - O:7 (C1-C4) | O:17 (J) | 1 | 0.167 | 2.935 | 1 | 0 | 0
# | 111 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35068_H4.0 |  | CVUAS 35068 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.955 | 2.439 | 1 | 0 | 0
| 111 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35068_H5.0 |  | CVUAS 35068 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.946 | 2.449 | 1 | 0 | 0
| 111 | 231212-1451-Loel Sal Vali Gr 7_CVUAS 35068_H6.0 |  | CVUAS 35068 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.94 | 2.645 | 1 | 0 | 0
| 111 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35068_H4.0 |  | CVUAS 35068 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.901 | 2.05 | 1 | 0 | 0
| 111 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35068_H5.0 |  | CVUAS 35068 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.888 | 2.013 | 1 | 0 | 0
| 111 | 231220-1151-BayP-Sal Vali Gr 7_CVUAS 35068_H6.0 |  | CVUAS 35068 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.91 | 1.837 | 1 | 0 | 0
# | 112 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11193_B7.0 |  | CVUAS 11193 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.774 | 3.63 | 0 | 1 | 0
| 112 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11193_B8.0 |  | CVUAS 11193 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.787 | 2.477 | 1 | 0 | 0
| 112 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11193_B9.0 |  | CVUAS 11193 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.547 | 2.528 | 1 | 0 | 0
| 112 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11193_B7.0 |  | CVUAS 11193 | FALSE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.67 | 2.505 | 1 | 0 | 0
| 112 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11193_B8.0 |  | CVUAS 11193 | FALSE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.811 | 2.872 | 1 | 0 | 0
| 112 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11193_B9.0 |  | CVUAS 11193 | FALSE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.845 | 2.107 | 1 | 0 | 0
| 112 | 231222-1357-BayP-Sal Vali WDH_CVUAS 11193_C4.0 |  | CVUAS 11193 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.908 | 2.415 | 1 | 0 | 0
| 112 | 231222-1357-BayP-Sal Vali WDH_CVUAS 11193_C5.0 |  | CVUAS 11193 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.806 | 3.313 | 1 | 0 | 0
| 112 | 231222-1357-BayP-Sal Vali WDH_CVUAS 11193_C6.0 |  | CVUAS 11193 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.879 | 2.484 | 1 | 0 | 0
# | 113 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5883_4_G10.0 |  | CVUAS 5883.4 | TRUE | non - O:7 (C1-C4) | O:21 (L) | 1 | 0.694 | 3.179 | 1 | 0 | 0
| 113 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5883_4_G11.0 |  | CVUAS 5883.4 | TRUE | non - O:7 (C1-C4) | O:21 (L) | 1 | 0.786 | 3.131 | 1 | 0 | 0
| 113 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5883_4_G12.0 |  | CVUAS 5883.4 | TRUE | non - O:7 (C1-C4) | O:21 (L) | 1 | 0.699 | 3.2 | 1 | 0 | 0
| 113 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5883_4_G10.0 |  | CVUAS 5883.4 | TRUE | non - O:7 (C1-C4) | O:21 (L) | 1 | 0.787 | 2.267 | 1 | 0 | 0
| 113 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5883_4_G11.0 |  | CVUAS 5883.4 | TRUE | non - O:7 (C1-C4) | O:21 (L) | 1 | 0.7 | 2.371 | 1 | 0 | 0
| 113 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5883_4_G12.0 |  | CVUAS 5883.4 | TRUE | non - O:7 (C1-C4) | O:21 (L) | 1 | 0.766 | 2.372 | 1 | 0 | 0
# | 114 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 8148_2_F7.0 |  | CVUAS 8148.2 | TRUE | non - O:7 (C1-C4) | O:21 (L) | 1 | 0.507 | 2.969 | 1 | 0 | 0
| 114 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 8148_2_F8.0 |  | CVUAS 8148.2 | TRUE | non - O:7 (C1-C4) | O:21 (L) | 1 | 0.605 | 2.977 | 1 | 0 | 0
| 114 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 8148_2_F9.0 |  | CVUAS 8148.2 | TRUE | non - O:7 (C1-C4) | O:21 (L) | 1 | 0.545 | 3.028 | 1 | 0 | 0
| 114 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 8148_2_F7.0 |  | CVUAS 8148.2 | TRUE | non - O:7 (C1-C4) | O:21 (L) | 1 | 0.311 | 2.208 | 1 | 0 | 0
| 114 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 8148_2_F8.0 |  | CVUAS 8148.2 | TRUE | non - O:7 (C1-C4) | O:21 (L) | 1 | 0.4 | 2.273 | 1 | 0 | 0
| 114 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 8148_2_F9.0 |  | CVUAS 8148.2 | TRUE | non - O:7 (C1-C4) | O:21 (L) | 1 | 0.351 | 2.189 | 1 | 0 | 0
# | 115 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 6882_D4.0 |  | CVUAS 6882 | TRUE | non - O:7 (C1-C4) | O:28 (M) | 1 | 0.867 | 2.764 | 1 | 0 | 0
| 115 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 6882_D5.0 |  | CVUAS 6882 | TRUE | non - O:7 (C1-C4) | O:28 (M) | 1 | 0.881 | 2.945 | 1 | 0 | 0
| 115 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 6882_D6.0 |  | CVUAS 6882 | TRUE | non - O:7 (C1-C4) | O:28 (M) | 1 | 0.883 | 2.815 | 1 | 0 | 0
| 115 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 6882_D4.0 |  | CVUAS 6882 | TRUE | non - O:7 (C1-C4) | O:28 (M) | 1 | 0.355 | 2.576 | 1 | 0 | 0
| 115 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 6882_D5.0 |  | CVUAS 6882 | TRUE | non - O:7 (C1-C4) | O:28 (M) | 1 | 0.436 | 2.651 | 1 | 0 | 0
| 115 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 6882_D6.0 |  | CVUAS 6882 | FALSE | non - O:7 (C1-C4) | O:28 (M) | 1 | 0.179 | 2.742 | 1 | 0 | 0
# | 116 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 732_2_A7.0 |  | CVUAS 732.2 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.119 | 1.884 | 1 | 0 | 0
| 116 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 732_2_A8.0 |  | CVUAS 732.2 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.152 | 1.757 | 1 | 0 | 0
| 116 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 732_2_A9.0 |  | CVUAS 732.2 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.222 | 1.658 | 1 | 0 | 0
| 116 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 732_2_A7.0 |  | CVUAS 732.2 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.173 | 1.896 | 1 | 0 | 0
| 116 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 732_2_A8.0 |  | CVUAS 732.2 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.17 | 1.775 | 1 | 0 | 0
| 116 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 732_2_A9.0 |  | CVUAS 732.2 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.16 | 1.753 | 1 | 0 | 0
# | 117 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9334_C10.0 |  | CVUAS 9334 | TRUE | non - O:7 (C1-C4) | O:30 (N) | 1 | 0.953 | 2.528 | 1 | 0 | 0
| 117 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9334_C11.0 |  | CVUAS 9334 | TRUE | non - O:7 (C1-C4) | O:30 (N) | 1 | 0.958 | 2.403 | 1 | 0 | 0
| 117 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 9334_C12.0 |  | CVUAS 9334 | TRUE | non - O:7 (C1-C4) | O:30 (N) | 1 | 0.958 | 2.612 | 1 | 0 | 0
| 117 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9334_C10.0 |  | CVUAS 9334 | TRUE | non - O:7 (C1-C4) | O:30 (N) | 1 | 0.876 | 2.544 | 1 | 0 | 0
| 117 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9334_C11.0 |  | CVUAS 9334 | TRUE | non - O:7 (C1-C4) | O:30 (N) | 1 | 0.92 | 2.73 | 1 | 0 | 0
| 117 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 9334_C12.0 |  | CVUAS 9334 | TRUE | non - O:7 (C1-C4) | O:30 (N) | 1 | 0.947 | 2.509 | 1 | 0 | 0
# | 118 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 1718_H7.0 |  | CVUAS 1718 | TRUE | non - O:7 (C1-C4) | O:30 (N) | 1 | 0.966 | 2.138 | 1 | 0 | 0
| 118 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 1718_H8.0 |  | CVUAS 1718 | TRUE | non - O:7 (C1-C4) | O:30 (N) | 1 | 0.969 | 2.253 | 1 | 0 | 0
| 118 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 1718_H9.0 |  | CVUAS 1718 | TRUE | non - O:7 (C1-C4) | O:30 (N) | 1 | 0.955 | 2.41 | 1 | 0 | 0
| 118 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 1718_H7.0 |  | CVUAS 1718 | TRUE | non - O:7 (C1-C4) | O:30 (N) | 1 | 0.964 | 2.286 | 1 | 0 | 0
| 118 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 1718_H8.0 |  | CVUAS 1718 | TRUE | non - O:7 (C1-C4) | O:30 (N) | 1 | 0.968 | 2.365 | 1 | 0 | 0
| 118 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 1718_H9.0 |  | CVUAS 1718 | TRUE | non - O:7 (C1-C4) | O:30 (N) | 1 | 0.968 | 2.156 | 1 | 0 | 0
# | 119 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 1473_2_G7.0 |  | CVUAS 1473.2 | TRUE | non - O:7 (C1-C4) | O:35 (O) | 1 | 0.886 | 2.078 | 1 | 0 | 0
| 119 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 1473_2_G8.0 |  | CVUAS 1473.2 | FALSE | non - O:7 (C1-C4) | O:35 (O) | 1 | 0.771 | 1.638 | 1 | 0 | 0
| 119 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 1473_2_G9.0 |  | CVUAS 1473.2 | TRUE | non - O:7 (C1-C4) | O:35 (O) | 1 | 0.917 | 2.392 | 1 | 0 | 0
| 119 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 1473_2_G7.0 |  | CVUAS 1473.2 | TRUE | non - O:7 (C1-C4) | O:35 (O) | 1 | 0.833 | 2.038 | 1 | 0 | 0
| 119 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 1473_2_G8.0 |  | CVUAS 1473.2 | TRUE | non - O:7 (C1-C4) | O:35 (O) | 1 | 0.857 | 2.054 | 1 | 0 | 0
| 119 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 1473_2_G9.0 |  | CVUAS 1473.2 | TRUE | non - O:7 (C1-C4) | O:35 (O) | 1 | 0.814 | 2.267 | 1 | 0 | 0
# | 120 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 3881 _H10.0 |  | CVUAS 3881 | TRUE | non - O:7 (C1-C4) | O:35 (O) | 1 | 0.8 | 2.241 | 1 | 0 | 0
| 120 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 3881 _H11.0 |  | CVUAS 3881 | TRUE | non - O:7 (C1-C4) | O:35 (O) | 1 | 0.827 | 2.265 | 1 | 0 | 0
| 120 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 3881 _H12.0 |  | CVUAS 3881 | TRUE | non - O:7 (C1-C4) | O:35 (O) | 1 | 0.945 | 2.409 | 1 | 0 | 0
| 120 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 3881 _H10.0 |  | CVUAS 3881 | TRUE | non - O:7 (C1-C4) | O:35 (O) | 1 | 0.923 | 2.274 | 1 | 0 | 0
| 120 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 3881 _H11.0 |  | CVUAS 3881 | TRUE | non - O:7 (C1-C4) | O:35 (O) | 1 | 0.948 | 2.254 | 1 | 0 | 0
| 120 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 3881 _H12.0 |  | CVUAS 3881 | TRUE | non - O:7 (C1-C4) | O:35 (O) | 1 | 0.944 | 2.098 | 1 | 0 | 0
# | 121 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4672_E4.0 |  | CVUAS 4672 | TRUE | non - O:7 (C1-C4) | O:38 (P) | 1 | 0.992 | 1.969 | 1 | 0 | 0
| 121 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4672_E5.0 |  | CVUAS 4672 | TRUE | non - O:7 (C1-C4) | O:38 (P) | 1 | 0.99 | 2.184 | 1 | 0 | 0
| 121 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4672_E6.0 |  | CVUAS 4672 | TRUE | non - O:7 (C1-C4) | O:38 (P) | 1 | 0.991 | 2.271 | 1 | 0 | 0
| 121 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4672_E4.0 |  | CVUAS 4672 | TRUE | non - O:7 (C1-C4) | O:38 (P) | 1 | 0.971 | 2.086 | 1 | 0 | 0
| 121 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4672_E5.0 |  | CVUAS 4672 | TRUE | non - O:7 (C1-C4) | O:38 (P) | 1 | 0.985 | 2.4 | 1 | 0 | 0
| 121 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4672_E6.0 |  | CVUAS 4672 | TRUE | non - O:7 (C1-C4) | O:38 (P) | 1 | 0.986 | 2.448 | 1 | 0 | 0
# | 122 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7149_G4.0 |  | CVUAS 7149 | TRUE | non - O:7 (C1-C4) | O:38 (P) | 1 | 0.994 | 1.708 | 1 | 0 | 0
| 122 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7149_G5.0 |  | CVUAS 7149 | TRUE | non - O:7 (C1-C4) | O:38 (P) | 1 | 0.994 | 1.903 | 1 | 0 | 0
| 122 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 7149_G6.0 |  | CVUAS 7149 | TRUE | non - O:7 (C1-C4) | O:38 (P) | 1 | 0.994 | 1.861 | 1 | 0 | 0
| 122 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7149_G4.0 |  | CVUAS 7149 | FALSE | non - O:7 (C1-C4) | O:38 (P) | 1 | 0.971 | 1.89 | 1 | 0 | 0
| 122 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7149_G5.0 |  | CVUAS 7149 | TRUE | non - O:7 (C1-C4) | O:38 (P) | 1 | 0.99 | 1.865 | 1 | 0 | 0
| 122 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 7149_G6.0 |  | CVUAS 7149 | TRUE | non - O:7 (C1-C4) | O:38 (P) | 1 | 0.988 | 2.276 | 1 | 0 | 0
# | 123 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 2189_A10.0 |  | CVUAS 2189 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.446 | 3.139 | 1 | 0 | 0
| 123 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 2189_A11.0 |  | CVUAS 2189 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.329 | 3.279 | 1 | 0 | 0
| 123 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 2189_A12.0 |  | CVUAS 2189 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.517 | 2.985 | 1 | 0 | 0
| 123 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 2189_A10.0 |  | CVUAS 2189 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.509 | 3.105 | 1 | 0 | 0
| 123 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 2189_A11.0 |  | CVUAS 2189 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.546 | 2.91 | 1 | 0 | 0
| 123 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 2189_A12.0 |  | CVUAS 2189 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.484 | 2.907 | 1 | 0 | 0
# | 124 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 8418_H7.0 |  | CVUAS 8418 | TRUE | non - O:7 (C1-C4) | O:6,14 (H) | 1 | 0.235 | 3.113 | 1 | 0 | 0
| 124 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 8418_H8.0 |  | CVUAS 8418 | TRUE | non - O:7 (C1-C4) | O:6,14 (H) | 1 | 0.22 | 3.167 | 1 | 0 | 0
| 124 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 8418_H9.0 |  | CVUAS 8418 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.162 | 2.68 | 1 | 0 | 0
| 124 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 8418_H7.0 |  | CVUAS 8418 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.4 | 1.956 | 1 | 0 | 0
| 124 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 8418_H8.0 |  | CVUAS 8418 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.311 | 2.069 | 1 | 0 | 0
| 124 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 8418_H9.0 |  | CVUAS 8418 | TRUE | non - O:7 (C1-C4) | O:56 | 1 | 0.38 | 1.883 | 1 | 0 | 0
# | 125 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4907_A7.0 |  | CVUAS 4907 | TRUE | non - O:7 (C1-C4) | O:41 (S) | 1 | 0.539 | 2.602 | 1 | 0 | 0
| 125 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4907_A8.0 |  | CVUAS 4907 | TRUE | non - O:7 (C1-C4) | O:41 (S) | 1 | 0.562 | 2.613 | 1 | 0 | 0
| 125 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4907_A9.0 |  | CVUAS 4907 | TRUE | non - O:7 (C1-C4) | O:41 (S) | 1 | 0.49 | 2.588 | 1 | 0 | 0
| 125 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4907_A7.0 |  | CVUAS 4907 | TRUE | non - O:7 (C1-C4) | O:41 (S) | 1 | 0.537 | 2.212 | 1 | 0 | 0
| 125 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4907_A8.0 |  | CVUAS 4907 | TRUE | non - O:7 (C1-C4) | O:41 (S) | 1 | 0.551 | 2.212 | 1 | 0 | 0
| 125 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4907_A9.0 |  | CVUAS 4907 | TRUE | non - O:7 (C1-C4) | O:41 (S) | 1 | 0.517 | 2.277 | 1 | 0 | 0
# | 126 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4976_3_C7.0 |  | CVUAS 4976.3 | TRUE | non - O:7 (C1-C4) | O:41 (S) | 1 | 0.574 | 2.379 | 1 | 0 | 0
| 126 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4976_3_C8.0 |  | CVUAS 4976.3 | TRUE | non - O:7 (C1-C4) | O:41 (S) | 1 | 0.458 | 2.037 | 1 | 0 | 0
| 126 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4976_3_C9.0 |  | CVUAS 4976.3 | TRUE | non - O:7 (C1-C4) | O:41 (S) | 1 | 0.519 | 2.067 | 1 | 0 | 0
| 126 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4976_3_C7.0 |  | CVUAS 4976.3 | TRUE | non - O:7 (C1-C4) | O:41 (S) | 1 | 0.577 | 2.287 | 1 | 0 | 0
| 126 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4976_3_C8.0 |  | CVUAS 4976.3 | TRUE | non - O:7 (C1-C4) | O:41 (S) | 1 | 0.539 | 2.13 | 1 | 0 | 0
| 126 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4976_3_C9.0 |  | CVUAS 4976.3 | TRUE | non - O:7 (C1-C4) | O:41 (S) | 1 | 0.522 | 2.118 | 1 | 0 | 0
# | 127 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5529_3_B10.0 |  | CVUAS 5529.3 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.434 | 2.114 | 1 | 0 | 0
| 127 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5529_3_B11.0 |  | CVUAS 5529.3 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.796 | 2.144 | 1 | 0 | 0
| 127 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5529_3_B12.0 |  | CVUAS 5529.3 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.802 | 2.217 | 1 | 0 | 0
| 127 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5529_3_B10.0 |  | CVUAS 5529.3 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.152 | 2.171 | 1 | 0 | 0
| 127 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5529_3_B11.0 |  | CVUAS 5529.3 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.167 | 2.47 | 1 | 0 | 0
| 127 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5529_3_B12.0 |  | CVUAS 5529.3 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.222 | 2.518 | 1 | 0 | 0
# | 128 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32015_B7.0 |  | CVUAS 32015 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.98 | 2.942 | 1 | 0 | 0
| 128 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32015_B8.0 |  | CVUAS 32015 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.982 | 2.973 | 1 | 0 | 0
| 128 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32015_B9.0 |  | CVUAS 32015 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.981 | 2.89 | 1 | 0 | 0
| 128 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32015_B7.0 |  | CVUAS 32015 | TRUE | non - O:7 (C1-C4) | O:16 (I) | 1 | 0.203 | 4.879 | 0 | 1 | 0
| 128 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32015_B8.0 |  | CVUAS 32015 | TRUE | non - O:7 (C1-C4) | O:16 (I) | 1 | 0.207 | 4.989 | 0 | 1 | 0
| 128 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32015_B9.0 |  | CVUAS 32015 | TRUE | non - O:7 (C1-C4) | O:16 (I) | 1 | 0.216 | 4.386 | 0 | 1 | 0
# | 129 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5493_A10.0 |  | CVUAS 5493 | TRUE | non - O:7 (C1-C4) | O:42 (T) | 1 | 0.501 | 2.497 | 1 | 0 | 0
| 129 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5493_A11.0 |  | CVUAS 5493 | TRUE | non - O:7 (C1-C4) | O:42 (T) | 1 | 0.62 | 2.368 | 1 | 0 | 0
| 129 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 5493_A12.0 |  | CVUAS 5493 | TRUE | non - O:7 (C1-C4) | O:42 (T) | 1 | 0.75 | 2.335 | 1 | 0 | 0
| 129 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5493_A10.0 |  | CVUAS 5493 | TRUE | non - O:7 (C1-C4) | O:42 (T) | 1 | 0.76 | 2.272 | 1 | 0 | 0
| 129 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5493_A11.0 |  | CVUAS 5493 | TRUE | non - O:7 (C1-C4) | O:42 (T) | 1 | 0.663 | 2.503 | 1 | 0 | 0
| 129 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 5493_A12.0 |  | CVUAS 5493 | TRUE | non - O:7 (C1-C4) | O:42 (T) | 1 | 0.834 | 2.334 | 1 | 0 | 0
# | 130 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32483_G7.0 |  | CVUAS 32483 | FALSE | non - O:7 (C1-C4) | O:42 (T) | 1 | 0.983 | 2.238 | 1 | 0 | 0
| 130 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32483_G8.0 |  | CVUAS 32483 | TRUE | non - O:7 (C1-C4) | O:42 (T) | 1 | 0.977 | 2.372 | 1 | 0 | 0
| 130 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32483_G9.0 |  | CVUAS 32483 | TRUE | non - O:7 (C1-C4) | O:42 (T) | 1 | 0.982 | 2.158 | 1 | 0 | 0
| 130 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32483_G7.0 |  | CVUAS 32483 | TRUE | non - O:7 (C1-C4) | O:42 (T) | 1 | 0.979 | 2.422 | 1 | 0 | 0
| 130 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32483_G8.0 |  | CVUAS 32483 | TRUE | non - O:7 (C1-C4) | O:42 (T) | 1 | 0.98 | 2.422 | 1 | 0 | 0
| 130 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32483_G9.0 |  | CVUAS 32483 | TRUE | non - O:7 (C1-C4) | O:42 (T) | 1 | 0.983 | 2.454 | 1 | 0 | 0
# | 131 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32796_C10.0 |  | CVUAS 32796 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.318 | 2.483 | 1 | 0 | 0
| 131 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32796_C11.0 |  | CVUAS 32796 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.367 | 2.625 | 1 | 0 | 0
| 131 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32796_C12.0 |  | CVUAS 32796 | TRUE | non - O:7 (C1-C4) | O:43 (U) | 1 | 0.293 | 2.808 | 1 | 0 | 0
| 131 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32796_C10.0 |  | CVUAS 32796 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.36 | 2.721 | 1 | 0 | 0
| 131 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32796_C11.0 |  | CVUAS 32796 | TRUE | non - O:7 (C1-C4) | O:13 (G) | 1 | 0.358 | 2.728 | 1 | 0 | 0
| 131 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32796_C12.0 |  | CVUAS 32796 | TRUE | non - O:7 (C1-C4) | O:43 (U) | 1 | 0.277 | 2.887 | 1 | 0 | 0
# | 132 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 456_3_G4.0 |  | CVUAS 456.3 | TRUE | non - O:7 (C1-C4) | O:43 (U) | 1 | 0.518 | 2.118 | 1 | 0 | 0
| 132 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 456_3_G5.0 |  | CVUAS 456.3 | TRUE | non - O:7 (C1-C4) | O:43 (U) | 1 | 0.497 | 2.019 | 1 | 0 | 0
| 132 | 231220-1435-Loel Sal Vali Gr 1_CVUAS 456_3_G6.0 |  | CVUAS 456.3 | TRUE | non - O:7 (C1-C4) | O:43 (U) | 1 | 0.509 | 2.289 | 1 | 0 | 0
| 132 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 456_3_G4.0 |  | CVUAS 456.3 | TRUE | non - O:7 (C1-C4) | O:43 (U) | 1 | 0.584 | 2.132 | 1 | 0 | 0
| 132 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 456_3_G5.0 |  | CVUAS 456.3 | TRUE | non - O:7 (C1-C4) | O:43 (U) | 1 | 0.539 | 2.109 | 1 | 0 | 0
| 132 | 231221-1306-BayP-Sal Vali Gr 1_CVUAS 456_3_G6.0 |  | CVUAS 456.3 | TRUE | non - O:7 (C1-C4) | O:43 (U) | 1 | 0.601 | 2.223 | 1 | 0 | 0
# | 133 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11433_E7.0 |  | CVUAS 11433 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.764 | 2.806 | 1 | 0 | 0
| 133 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11433_E8.0 |  | CVUAS 11433 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.792 | 2.841 | 1 | 0 | 0
| 133 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 11433_E9.0 |  | CVUAS 11433 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.673 | 2.985 | 1 | 0 | 0
| 133 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11433_E7.0 |  | CVUAS 11433 | FALSE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.751 | 2.543 | 1 | 0 | 0
| 133 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11433_E8.0 |  | CVUAS 11433 | FALSE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.64 | 2.431 | 1 | 0 | 0
| 133 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 11433_E9.0 |  | CVUAS 11433 | FALSE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.75 | 2.174 | 1 | 0 | 0
| 133 | 231221-1432-BayP-Sal Vali WDH_CVUAS 11433_H4.0 |  | CVUAS 11433 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.736 | 2.55 | 1 | 0 | 0
| 133 | 231221-1432-BayP-Sal Vali WDH_CVUAS 11433_H5.0 |  | CVUAS 11433 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.714 | 2.681 | 1 | 0 | 0
| 133 | 231221-1432-BayP-Sal Vali WDH_CVUAS 11433_H6.0 |  | CVUAS 11433 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.668 | 2.559 | 1 | 0 | 0
# | 134 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32016_C7.0 |  | CVUAS 32016 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.776 | 2.242 | 1 | 0 | 0
| 134 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32016_C8.0 |  | CVUAS 32016 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.659 | 2.293 | 1 | 0 | 0
| 134 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 32016_C9.0 |  | CVUAS 32016 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.752 | 2.187 | 1 | 0 | 0
| 134 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32016_C7.0 |  | CVUAS 32016 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.792 | 2.319 | 1 | 0 | 0
| 134 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32016_C8.0 |  | CVUAS 32016 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.758 | 2.427 | 1 | 0 | 0
| 134 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 32016_C9.0 |  | CVUAS 32016 | TRUE | non - O:7 (C1-C4) | O:44 (V) | 1 | 0.742 | 2.415 | 1 | 0 | 0
# | 135 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33139_C4.0 |  | CVUAS 33139 | TRUE | non - O:7 (C1-C4) | O:45 (W) | 1 | 0.553 | 2.018 | 1 | 0 | 0
| 135 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33139_C5.0 |  | CVUAS 33139 | TRUE | non - O:7 (C1-C4) | O:45 (W) | 1 | 0.543 | 2.471 | 1 | 0 | 0
| 135 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33139_C6.0 |  | CVUAS 33139 | TRUE | non - O:7 (C1-C4) | O:45 (W) | 1 | 0.504 | 2.445 | 1 | 0 | 0
| 135 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33139_C4.0 |  | CVUAS 33139 | TRUE | non - O:7 (C1-C4) | O:45 (W) | 1 | 0.579 | 1.562 | 1 | 0 | 0
| 135 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33139_C5.0 |  | CVUAS 33139 | TRUE | non - O:7 (C1-C4) | O:45 (W) | 1 | 0.543 | 1.954 | 1 | 0 | 0
| 135 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33139_C6.0 |  | CVUAS 33139 | TRUE | non - O:7 (C1-C4) | O:45 (W) | 1 | 0.593 | 2.063 | 1 | 0 | 0
# | 136 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 10812_3_F4.0 |  | CVUAS 10812.3 | TRUE | non - O:7 (C1-C4) | O:45 (W) | 1 | 0.94 | 2.523 | 1 | 0 | 0
| 136 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 10812_3_F5.0 |  | CVUAS 10812.3 | TRUE | non - O:7 (C1-C4) | O:45 (W) | 1 | 0.95 | 2.472 | 1 | 0 | 0
| 136 | 231214-1205-Loel Sal Vali Gr 4_CVUAS 10812_3_F6.0 |  | CVUAS 10812.3 | TRUE | non - O:7 (C1-C4) | O:45 (W) | 1 | 0.957 | 2.351 | 1 | 0 | 0
| 136 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 10812_3_F4.0 |  | CVUAS 10812.3 | TRUE | non - O:7 (C1-C4) | O:45 (W) | 1 | 0.969 | 1.981 | 1 | 0 | 0
| 136 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 10812_3_F5.0 |  | CVUAS 10812.3 | FALSE | non - O:7 (C1-C4) | O:45 (W) | 1 | 0.948 | 2.162 | 1 | 0 | 0
| 136 | 231215-1202-BayP-Sal Vali Gr 4_CVUAS 10812_3_F6.0 |  | CVUAS 10812.3 | FALSE | non - O:7 (C1-C4) | O:45 (W) | 1 | 0.976 | 2.097 | 1 | 0 | 0
# | 137 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 8118_2_B10.0 |  | CVUAS 8118.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.205 | 4.941 | 0 | 1 | 0
| 137 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 8118_2_B11.0 |  | CVUAS 8118.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.21 | 4.727 | 0 | 1 | 0
| 137 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 8118_2_B12.0 |  | CVUAS 8118.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.228 | 4.819 | 0 | 1 | 0
| 137 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 8118_2_F10.0 |  | CVUAS 8118.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.182 | 4.626 | 0 | 1 | 0
| 137 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 8118_2_F11.0 |  | CVUAS 8118.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.222 | 4.748 | 0 | 1 | 0
| 137 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 8118_2_F12.0 |  | CVUAS 8118.2 | TRUE | non - O:7 (C1-C4) | O:4 (B) | 1 | 0.225 | 4.807 | 0 | 1 | 0
# | 138 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33326_F4.0 |  | CVUAS 33326 | TRUE | non - O:7 (C1-C4) | O:47 (X) | 1 | 0.298 | 2.022 | 1 | 0 | 0
| 138 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33326_F5.0 |  | CVUAS 33326 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.184 | 2.392 | 1 | 0 | 0
| 138 | 231214-1417-Dyk-Sal Vali Gr 6_CVUAS 33326_F6.0 |  | CVUAS 33326 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.202 | 2.611 | 1 | 0 | 0
| 138 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33326_F4.0 |  | CVUAS 33326 | TRUE | non - O:7 (C1-C4) | O:47 (X) | 1 | 0.576 | 1.715 | 1 | 0 | 0
| 138 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33326_F5.0 |  | CVUAS 33326 | TRUE | non - O:7 (C1-C4) | O:65 | 1 | 0.456 | 2.538 | 1 | 0 | 0
| 138 | 231219-1427-BayP-Sal Vali Gr 6_CVUAS 33326_F6.0 |  | CVUAS 33326 | TRUE | non - O:7 (C1-C4) | O:47 (X) | 1 | 0.722 | 1.577 | 1 | 0 | 0
# | 139 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 30643_2_C10.0 |  | CVUAS 30643.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.993 | 1.792 | 1 | 0 | 0
| 139 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 30643_2_C11.0 |  | CVUAS 30643.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.994 | 1.847 | 1 | 0 | 0
| 139 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 30643_2_C12.0 |  | CVUAS 30643.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.994 | 1.799 | 1 | 0 | 0
| 139 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 30643_2_G10.0 |  | CVUAS 30643.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.992 | 1.978 | 1 | 0 | 0
| 139 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 30643_2_G11.0 |  | CVUAS 30643.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.992 | 2.004 | 1 | 0 | 0
| 139 | 240131-1400-Loel-BayP-Sali Vali WDH_CVUAS 30643_2_G12.0 |  | CVUAS 30643.2 | TRUE | non - O:7 (C1-C4) | O:9 (D1) | 1 | 0.993 | 2.066 | 1 | 0 | 0
# | 140 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4970_3_B7.0 |  | CVUAS 4970.3 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.147 | 2.658 | 1 | 0 | 0
| 140 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4970_3_B8.0 |  | CVUAS 4970.3 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.132 | 2.521 | 1 | 0 | 0
| 140 | 231212-0934-Dyk-Sal Vali Gr2_CVUAS 4970_3_B9.0 |  | CVUAS 4970.3 | TRUE | non - O:7 (C1-C4) | O:18 (K) | 1 | 0.128 | 2.48 | 1 | 0 | 0
| 140 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4970_3_B7.0 |  | CVUAS 4970.3 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.745 | 2.23 | 1 | 0 | 0
| 140 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4970_3_B8.0 |  | CVUAS 4970.3 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.861 | 2.334 | 1 | 0 | 0
| 140 | 231219-1154-Loel-Sal Vali Gr2_CVUAS 4970_3_B9.0 |  | CVUAS 4970.3 | TRUE | non - O:7 (C1-C4) | Others | 1 | 0.509 | 2.384 | 1 | 0 | 0
# | 141 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 6899_2_E4.0 |  | CVUAS 6899.2 | TRUE | non - O:7 (C1-C4) | O:35 (O) | 1 | 0.201 | 2.662 | 1 | 0 | 0
| 141 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 6899_2_E5.0 |  | CVUAS 6899.2 | TRUE | non - O:7 (C1-C4) | O:43 (U) | 1 | 0.159 | 2.445 | 1 | 0 | 0
| 141 | 231213-0850-Dyk-Sal Vali Gr3_CVUAS 6899_2_E6.0 |  | CVUAS 6899.2 | TRUE | non - O:7 (C1-C4) | O:43 (U) | 1 | 0.156 | 2.589 | 1 | 0 | 0
| 141 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 6899_2_E4.0 |  | CVUAS 6899.2 | TRUE | non - O:7 (C1-C4) | O:43 (U) | 1 | 0.2 | 2.241 | 1 | 0 | 0
| 141 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 6899_2_E5.0 |  | CVUAS 6899.2 | TRUE | non - O:7 (C1-C4) | O:43 (U) | 1 | 0.244 | 2.363 | 1 | 0 | 0
| 141 | 231214-0940-BayP-Sal Vali Gr3_CVUAS 6899_2_E6.0 |  | CVUAS 6899.2 | FALSE | non - O:7 (C1-C4) | O:45 (W) | 1 | 0.207 | 2.016 | 1 | 0 | 0
# | 142 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 30646_2_A4.0 |  | CVUAS 30646.2 | TRUE | non - O:7 (C1-C4) | O:58 | 1 | 0.773 | 2.376 | 1 | 0 | 0
| 142 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 30646_2_A5.0 |  | CVUAS 30646.2 | TRUE | non - O:7 (C1-C4) | O:58 | 1 | 0.765 | 2.37 | 1 | 0 | 0
| 142 | 231213-1218-Loel Sal Vali Gr 5_CVUAS 30646_2_A6.0 |  | CVUAS 30646.2 | TRUE | non - O:7 (C1-C4) | O:58 | 1 | 0.787 | 2.415 | 1 | 0 | 0
| 142 | 231215-0930-Dyk-Sal Vali Gr 5_CVUAS 30646_2_A4.0 |  | CVUAS 30646.2 | TRUE | non - O:7 (C1-C4) | O:58 | 1 | 0.818 | 1.875 | 1 | 0 | 0

## Report
Targeted Identification
following the guidelines for validating species identifications using MALDI-TOF-MS pursuant to §64 of the German Food and Feed code (LFGB), coordinated by the Federal Office of Consumer Protection and Food Safety, Berlin, Germany (BVL) 28 Oct 2022
Salmonella O-group C1
Version: | 1
generated by | OBH | checked by | Dyk
generated on | 2024-02-08 | checked on | 2024-02-09
Application | Identification of Salmonella O-groups
Environment | Salmonella O:7 (C1-C4) and non - O:7 (C1-C4)
Applied database / software version |  | Bruker classifier Salmonella O-groups v3 | IR Biotyper v4.0 01/2023
Validation set
number of target parameter isolates: 25 |  | number of non-target parameter isolates: 142
number of target parameter spectra: 150 |  | number of non-target parameter spectra: 889
cultivation conditions: COL-SB BD 37°C 24h aerobic
filtration method: preparation acc. to producer instructions (Instruction for use IR Biotyper Kit, Revision B (03/2021))
Only identified samples are used for determination of the inclusivity and exclusivity.
True-positive: outlier score </=3,5 AND target parameter classification = label |  | True negative: outlier score </=3,5 AND non-target parameter classification <> target parameter
False negative: true-positive criteria aren't fullfilled |  | False positive: true-negative criteria aren't fullfilled
Validation result | Target parameter: | Salmonella O-group C1 |  | target | fulfilled
Total number of spectra | 150
of which with ID result | 148 | rate of identified spectra | 98.67% | 90% | yes | 1
Total number of isolates | 25 |  |  | 20 | yes | 1
With respect to identified spectra: |  | Inklusivity
Positive conformity (PA?) | 147 | True positive rate | 99.32% | 95% | yes | 1
Negative deviation (ND) | 1 | False negative rate | 0.6757% | 1% | yes | 1
Non-target parameter
Total number of spectra | 889
of which with ID result | 854 | rate of identified spectra | 96.06% | 90% | yes | 1
Total number of isolates | 142 |  |  | 30 | yes | 1
With respect to identified spectra: |  | Exklusivity
Positive conformity (PA?) | 854 | True negative rate | 100% | 99% | yes | 1
Negative deviation (ND) | 0 | False-Positive Rate | 0% | 1% | yes | 1
Out of 148 identified target parameter spectra, 147 (=99.3%) were classified correctly (inclusivity). 1 (=0.7%) of identified target parameter spectra were allocated incorrectly to a different serogroup.
Out of 854 identified non-target parameter spectra, 854 (=100%) were counted correctly as negative (exclusivity). 0 (=0%) of all non-target parameter spectra were misclassified as the target parameter.
Criteria for the | targeted identification | are thus | fulfilled
Comment:
Release for site | for department | on (date) | by (signature)
CVUA S | T, MT, D | 2024-02-09 | RAU
The procedure is suitable for the intended purpose and is thereby released for usage.
Analytical parameter added to LIMS:
A signed Version of this validation report is kept within the generating lab.
yes
no